System Requirements Specification:

Plantation Productions Open Source/Open Hardware Data Acquisition System

Table of Contents

1	Introduction	4

1.1	System Purpose	4

1.2	Scope	4

1.3	Intended Audience	5

1.4	Definitions, Acronyms and Abbreviations	5

1.4.1	Definitions	5

1.4.2	Acronyms and Abbreviations	5

1.5	Existing Solutions	6

1.6	Argument for an Open-Source/Open-Hardware Solution	6

1.7	Overview/Contents of the Document	7

1.8	Document Conventions	7

2	General System Description	8

2.1	Design Basis	8

2.2	System Context	8

2.2.1	Digital Data Acquisition Control and Communications (CPU Interface)	8

2.2.2	Digital Inputs	9

2.2.3	Optically Isolated Digital Inputs	9

2.2.4	Digital Outputs	9

2.2.5	Analog Inputs	9

2.2.6	Analog Input Conditioning	9

2.2.7	Analog Outputs	10

2.2.8	Analog Output Conditioning	10

2.3	I/O Scanning	11

2.4	Analog Conditioning Adjustment	11

2.5	Host Communication	12

2.6	System Regulatory Issues	12

2.7	User Characteristics	12

3	System Requirements	14

3.1	Physical Requirements	14

3.1.1	Module Width	14

3.1.2	Module Length	14

3.1.3	Mounting Holes	14

3.2	Electromechanical Requirements	14

3.2.1	Circuit Board Layers	14

3.2.2	Circuit Board Top Layer	14

3.2.3	Circuit Board Second (Interior) Layer	14

3.2.4	Circuit Board Third (Interior) Layer	14

3.2.5	Circuit Board Bottom Layer	14

3.2.6	Signal Trace Separation	14

3.2.7	Minimum Trace Width	15

3.2.8	Minimum Power Trace Width	15

3.2.9	Minimum High-Current Trace Width	15

3.3	Hardware Requirements	15

3.3.1	CPU Interface Board (DAQIF)	15

3.3.2	Parallel I/O Interface Board (PPDIO96)	19

3.3.3	12-Channel Optoisolator Parallel Input Board (PPOPTO-12)	25

3.3.4	16-Channel Solid-State Relay Output Board (PPSSR16)	27

3.3.5	12-Channel Mechanical Relay Output Board (PPRELAY12)	30

3.3.6	48-Channel Digital Output Board (PPDO-48)	33

3.3.7	16-Channel Analog Input/4-Channel Analog Output Board (PPAIO-16/4)	38

3.3.8	4-Channel Analog Conditioning Board (PPAC-4)	46

3.3.9	12-Channel Relay-based Digital I/O Board (PPRlyio-12)	48

3.3.10	PPAC420 4-20mA Analog Input Conditioning Board (PPAC420)	51

3.4	Software Requirements	51

3.4.1	DIS64 Emulation	51

3.4.2	PPDAQ Standard Control Software	53

3.5	System Security	99

3.6	System Operations	99

3.6.1	Human Factors	99

3.7	System Power	99

3.7.1	PPDAQ Input Power	99

3.7.2	PPAC4 Input Power	99

3.7.3	PPAIO16-4 Input Power	99

3.7.4	PPDIO96 Input Power	99

3.7.5	PPOPTO12 Input Power	99

3.7.6	PPRELAY Input Power	99

3.7.7	PPSSR16 Input Power	99

3.7.8	Netburner Input Power	100

3.7.9	Netburner Input Power -- USB	100

3.7.10	Ethernet Connection	100

3.7.11	RS-232 Serial Connection	100

3.7.12	USB Connection	100

4	Design Goals and Other Requirements	101

List of Figures

Figure 1 – New AFRRI Power Instrument Ranges	11

Figure 2 – AFRRI Control System Block Diagram	12

Figure 3 - Rod Drive	14

Figure 4 – Pneumatic Rod Drive	15

# Introduction

This document is the System Requirements Specification (SyRS) for the Plantation Productions' Open Source/Open Hardware digital data acquisition system. The system software and hardware is covered under the Creative Commons (CC BY 4.0) found here:

https://creativecommons.org/licenses/by/4.0/

For the purposes of attribution, all work must be attributed to "Randall Hyde, Plantation Productions, Inc., Copyright 2017"

The Plantation Productions' digital data acquisition system consists of a series of modules (each with their own set of documentation) that provides the following functionality:

- Interface to some CPU module  (typically a SBC such as a Netburner MOD54415, Raspberry Pi 3, or Teensy 3.2, although the requirements do not limit the interface to these particular computers)

- A watchdog timer module to ensure correct system operation

- A digital input module accepting TTL (5V) or dry contact inputs

- An opto-isolation digital input interface that connects with the digital input module

- A digital output module featuring mechanical relays

- A digital output module featuring DC solid-state relays

- An analog I/O module featuring ADC and DAC components

- An analog input conditioning module that converts instrumentation analog levels (±10V) to voltage levels acceptable to the ADC module

- An analog input conditioning module that converts 4-20mA signals to voltage levels in the range -1.25V to +5V (0ma to 20ma)

- An analog output conditioning module that converts the DAC signals to instrument-level voltages (±10V)

This is not an inclusive list. Additional modules may be added to the system at a later date. Examples might include thermocouple or RTD input modules.

The CPU module will provide an interface to a general-purpose computer system (e.g., PC) via Ethernet, USB, or RS-232 serial; not all interfaces are required to support all three data transmission protocols. The system will employ the same communications protocol regardless of the physical interface.

The end result is a single well-documented and well-tested design that is applicable to the nuclear industry for data acquisition, output, and control.

## Revision History

| Revision | Date | Author | Description |
| --- | --- | --- | --- |
| V1.0 |  | Randall Hyde | Initial Document |
| V1.1 | 04/14/2019 | Randall Hyde | Removed mention of RESET line from PPDIO96 requirements. |

## System Purpose

The purpose of the system is to provide an open, flexible, digital data acquisition system that can be incorporated into nuclear reactor control systems.

## Scope

The DAQ hardware and software will be produced from this specification.

The objectives of the hardware and software development are to provide functions, status information, monitor and control hardware, communications, internal and self-test functions per the requirements that have been allocated to the DAQ system.

## Intended Audience

The intended audience of this specification is the engineering, product assurance and management personnel involved in Plantation Productions' DAQ hardware and software development and for development of the Software Design Description (SDD) document.

## Definitions, Acronyms and Abbreviations

### Definitions

The definitions used herein are consistent with IEEE 610.12-1990 with the following clarifications:

| Accuracy | The degree of agreement with the true value of the measured input, expressed as percent of reading for digital readouts. (ANSI N42.18-1980) |
| --- | --- |
| Anomaly | Anything observed in the documentation or operation of software that deviates from expectations. [Derived from IEEE Std 610.12-1990] |
| Catastrophic event | A catastrophic event is an event without warning from which recovery is impossible. Catastrophic events include hardware or software failures resulting in computation and processing errors The processor will halt or reset, based on a configuration item, after a catastrophic event. |
| Channel | The features and capabilities associated with a detector, a sensor, or a calculated group of information. |
| CPU | Central Processing Unit; specifically, a microcomputer chip. |
| Failsafe Condition | Failsafe condition is an actuated state of the TRIGA System as the result of a catastrophic failure such as loss of power, break of circuit or device failure catastrophic event |
| Handled Conditions | Conditions that the system is designed to handle and continue processing. These conditions include anomalies, faults and failures. |
| Hardware Requirements Specification | A specification that documents the hardware requirements (HRS) |
| Power Failure | Power failure is the condition when AC power is outside required limits or logic power is below a low limit. |
| Precision (per ANSI N42.18-80) | The degree of agreement of repeated measurements of the same input, expressed as percent deviation from the mean reading at 95% confidence level. |
| SBC | Single-board computer |
| Software Requirements Specification (SRS) | Documentation of the essential requirements (functions, performance, design constraints, and attributes) of the software and its external interfaces. [IEEE Std 610.12-1990] |
| System Requirements Specification (SyRS) | A structured collection of information that embodies the requirements of the system. [IEEE Std 1233-1998] A specification that documents the requirements to establish a design basis and the conceptual design for a system or subsystem. [GA-ESI] |

### Acronyms and Abbreviations

The abbreviations listed have the following meanings where used in this specification:

| AC | Alternating Current |
| --- | --- |
| ANSI | American National Standards Institute |
| CPU | Central Processing Unit |
| Cpm | Counts per Minute |
| DAQ | Data Acquisition (System) |
| DPDT | Double-Pole Double-Throw |
| FAT | Factory Acceptance Test |
| IEEE | Institute of Electrical and Electronic Engineers |
| LED | Light-Emitting Diode |
| MTBF | Mean Time between Failures |
| NEMA | National Electrical Manufacturers Association |
| NRC | Nuclear Regulatory Commission |
| RAM | Radiation Area Monitor |
| Rms | Root Mean Square |
| SyRS | System Requirements Specification |
| Vac | Voltage Alternating Current |
| Vdc | Voltage Direct Current |
| WDT | Watchdog Timer |

## Existing Solutions

Existing solutions have either been ad-hoc custom designs or have used common-off-the-shelf (COTS) devices. COTS devices have the following drawbacks:

- Support: COTS devices have a lifetime that is much shorter than a typical device installation at a nuclear reactor. Obtaining replacement parts becomes difficult after end-of-life occurs on the device.

- Proprietary: COTS devices often employ a proprietary design and/or proprietary firmware making it difficult to perform a safety analysis on such devices or ensure that the device meets system requirements.

- Communication protocol: many COTS devices use a proprietary communication protocol making it difficult to substitute a different device, should that become necessary (e.g., when a device reaches end-of-life).

Custom data acquisition designs also have drawbacks:

- Custom designs are very expensive.

- Custom designs require considerable engineering resources and time that may not be available.

- Custom designs require considerable cross-disciplinary support (e.g., software, hardware, and quality assurance); such resources may not be available at a given site.

- Custom designs are difficult to share with other sites, thus limiting the amortization of costs to a single reactor site.

## Argument for an Open-Source/Open-Hardware Solution

An open-source/open-hardware digital data acquisition solution offers several advantages over COTS and custom solutions. Some of the advantages include:

- Development costs are amortized across multiple installations (versus custom designs where the expense is applied to a single site).

- Once the design exists, the number of engineering resources is reduced as most of the low-level design work has already been done.

- Open designs have a much longer lifetime that COTS designs. As parts become obsolete an open design allows minimal re-engineering to replace just the parts that are obsolete without a huge impact on the overall design.

- Open designs use an open communication protocol and underlying hardware/software design. This makes is possible to perform system analysis to ensure that the system operates properly.

One other aspect of an open design is that it must be extensible. No single data acquisition system will be suitable for all reactor sites. The number of I/O channels (analog and digital) vary greatly from site to site. Although one could adopt a "one size fits all" solution where the design implements the maximum number of I/O channels handled at any site, this may present other problems (such as size or power restrictions). A better solution is to allow a system designer to build a system from a set of I/O modules and add additional modules for larger systems (or use fewer modules for smaller systems).

## Overview/Contents of the Document

The general description including product perspective, product functions, user characteristics and general constraints is included in Section 2.

The specific requirements including external interface requirements, system features, performance requirements, and other requirements are included in Section 3.  The system features are specified.  For each system feature, an introduction is provided and the details for each stimulus/response are provided.

The traceability (reverse traceability) to other documents (e.g., SRS, SDD, HRS, STC, and STD) is documented in a separate document (Excel spreadsheet).

Section 2 contains the general description of the DAQ hardware and software.

The general description including product perspective, product functions, user characteristics and general constraints is included in this section.

Section 3 contains the specific requirements for the DAQ hardware and software.

The specific requirements including external interface requirements, system features, performance requirements, and other requirements are included in this section.

Section 4 contains other requirements or design goals

These are items that are not directly hardware or software related (that is, that derived from software design elements).

## Document Conventions

All system requirement tags shall take the form:

<whitespace>	[DAQ_SYRS_xxx]

<whitespace>	[DAQ_SYRS_xxx.yy]

<whitespace>	[DAQ_SYRS_xxx.yy.zz]

etc.

where "xxx" is a three- or four-digit requirement number.

For SYRS requirement tags, should the need arise to insert a new SYRS requirement tag between two other values (e.g., add a requirement between DAQ_SYRS_030 and DAQ_SYRS_031) then a decimal fractional number with exactly two digits shall be appended to the SYRS tag number (e.g., DAQ_SYRS_031.05). Any number of decimal point suffixes can be added, if needed (e.g., DAQ_SYRS_030.05.02).

# General System Description

## Design Basis

The design basis for the DAQ System (or just "System") is to provide a system that can be verified to be correct as per regulatory requirements (e.g., NRC requirements).

## System Context

For the purposes of this SyRS, the General System Description is laid out as:

- Digital data acquisition control and communications (CPU Interface)

- Digital input

- Digital output

- Analog input

- Analog output

### Digital Data Acquisition Control and Communications (CPU Interface)

The CPU interface module shall have the following responsibilities:

- Reading data from the digital and analog input devices

- Writing data to the digital and analog output devices

- Refreshing the watchdog timer

- Communicating with a host computer via Ethernet, USB, or RS-232 serial communication

The data acquisition and output modules shall communicate with a CPU via I2C and SPI communication busses and eight digital I/O lines (all 5V logic). In theory, you can interface any CPU or SBC that provides these I/O capabilities with the data acquisition and output modules. In practice, many CPUs/SBCs require additional logic-level shifting (e.g., between 3.3V and 5V) and an I2C expansion module, so the DAQ System will normally use a CPU interface board as an interface between the SBC/CPU and the data acquisition modules.

The CPU interface board is also a good place to put the watchdog timer circuitry.

The current CPU Interface board supports the Netburner MOD54415 MOD-DEV-70 evaluation board module, the Raspberry Pi 3 (any 40-pin Raspberry Pi module should also work), or a Teensy 3.2 module (high-performance Arduino-type board). The Netburner MOD-DEV-70 is the preferred SBC as it provides a true RTOS and Ethernet support (The Raspberry Pi is not real-time and the Teensy doesn't [directly] support Ethernet).

The existing CPU interface board also provides RS-232 level-shifting circuitry for the Teensy and Raspberry Pi SBCs (the Netburner provides this capability on-board, and doesn't use the circuitry on the CPU interface board).

Finally, the CPU interface board provides headers to connect with the various I/O modules using standard ribbon cabling.

Note: The Netburner and Teensy modules provide device-type USB support. To the host computer, these USB devices look like (very high speed) serial ports. Therefore, the host computer software actually sees only two types of devices: Ethernet and Serial, even though the System communicates with the host via USB, Ethernet, or RS-232 serial. Although the Raspberry Pi 3 SBC provides USB ports, these are all USB-host ports and cannot directly connect to the host computer system.  At this time, this specification ignores any USB ports (host or device) found on the Raspberry Pi SBCs.

### Digital Inputs

The System shall provide a minimum of 96 digital (TTL/5V) digital inputs with an expansion to more than 400 digital inputs as an option. By adding a pull-up resistor to +5V (or enabling on-chip pullup resistors), these inputs should also be capable of reading dry contact switch inputs.

### Optically Isolated Digital Inputs

The System shall provide the ability to isolate groups of digital inputs via opto-isolation. The opto-isolation circuitry will sit between the input device(s) and the digital input circuitry. The opto-isolator units shall accept voltage levels in the range 0-15V (programmable by changing a resistor) or dry contact inputs. All digital inputs shall be isolated from one another as well as from the DAQ System digital inputs.

### Digital Outputs

The System shall provide a minimum of 12 digital outputs, expandable to at least 96 digital outputs. These digital outputs can be TTL/5V logic outputs, higher-current TTL/5V outputs (at least 50mA/output), mechanical relay outputs, solid-state relay outputs, or any combination of these three. Mechanical relay outputs will, at a minimum, support normally-open contacts; if possible, they should support normally-closed and normally-open contacts. Relays (solid-state or mechanical) should be fail-safe normally open. That is, when power is not applied or the System is not initialized, the relays should be in the "off" (open) position. This requirement does not apply to TTL outputs.  Note: the TTL output requirement can be realized by the digital input module if the digital inputs support a selectable "data direction" allowing the digital I/O bits to be programmed as either an input or an output bit. In such cases, the digital input module shall default to input on reset/power up.

### Analog Inputs

The System shall provide a minimum of 16 single-ended or 8 double-ended (differential) analog inputs. It should be expandable to a minimum of 64 single-ended or 32 double-ended analog inputs. The ADCs should provide a minimum of 12 bits of precision (signed, two's complement value), 16-bits is preferred. At the very least, the ADC inputs should be capable of handling 0-3.3V; any other analog inputs must be preprocessed by analog conditioning circuitry.

### Analog Input Conditioning

Separate (from the analog input) analog conditioning circuitry shall be capable of handling ±10V inputs and translating this to the voltage range acceptable to the ADC (e.g., 0-5V). The signal conditioning circuitry shall produce a differential output to be fed into the double-ended inputs on the ADC.

Separate analog signal conditioning circuitry shall be capable of translating a 4-20mA signal to the voltage range acceptable by the analog input conditioning boards.

### Analog Outputs

The System shall provide a minimum of 4 analog outputs. It should be expandable to a minimum of 32 analog outputs. The DACs should provide a minimum of 12 bits of precision (signed, two's complement value). At the very least, the DAC inputs should be capable of producing 0-5V; any other analog inputs must be preprocessed by analog output conditioning circuitry.

### Analog Output Conditioning

Separate (from the analog output) analog conditioning circuitry shall be capable of converting the 0-5V output from the DAC to a range more acceptable to the System. Specifically, the analog output conditioning circuitry shall produce ±5V or ±10V when fed a 0-5V signal (note that 2.5V on input shall produce 0V on output).

#### CPU Interface Board

The data acquisition and output modules shall communicate with a CPU via I2C and SPI communication busses and eight digital I/O lines (all 5V logic). In theory, you can interface any CPU or SBC that provides these I/O capabilities with the data acquisition and output modules. In practice, many CPUs/SBCs require additional logic-level shifting (e.g., between 3.3V and 5V) and an I2C expansion module, so the DAQ System will normally use a CPU interface board as an interface between the SBC/CPU and the data acquisition modules.

The CPU interface board is also a good place to put the watchdog timer circuitry.

The current CPU Interface board supports the Netburner MOD54415 MOD-DEV-70 evaluation board module, the Raspberry Pi 3 (any 40-pin Raspberry Pi module should also work), or a Teensy 3.2 module (high-performance Arduino-type board). The Netburner MOD-DEV-70 is the preferred SBC as it provides a true RTOS and Ethernet support (The Raspberry Pi is not real-time and the Teensy doesn't [directly] support Ethernet).

The existing CPU interface board also provides RS-232 level-shifting circuitry for the Teensy and Raspberry Pi SBCs (the Netburner provides this capability on-board, and doesn't use the circuitry on the CPU interface board).

Finally, the CPU interface board provides headers to connect with the various I/O modules using standard ribbon cabling.

## I/O Scanning

The System shall periodically scan all the input devices and save the data for use by the host computer. During the scanning period, it shall also output any data received from the host computer. Note that the saved output data shall be written to the output devices (analog and digital) even if those values haven't been changed by the host computer since the last output cycle.

The I/O scanning period shall be no less than 12.5 msec and no more than 50 msec.

The System shall maintain a "history list" of at least one second's data of input readings since the last I/O request from the host. The list is a FIFO list; if the list overflows before the host reads the data, then the oldest readings are forgotten. Whenever the host reads the data, the System clears the list except for the most recent reading. If the inputs change in this list prior to reading the data, the System shall return data values (on an input-by-input basis) using one of the following mechanisms:

Digital:

- Most recent reading in history list

- First reading in history list

- Vote: the returned value is the largest number of zeros or ones (if there is a tie, the most recent reading wins).

- Loser: the returned value is the lesser number of zeros or ones (if there is a tie, the most recent reading wins).

- Debounce: the returned value is the longest sequence (with a minimum length specified by the host computer for the input) of a single value in the list. If no such sequence exists, the System returns the previous debounced value. If two or more sequences exist, the System returns the value of the most recent sequence).

Analog:

- Most recent reading in history list

- First reading in history list

- Minimum: smallest value in list

- Maximum: largest value in last

- Median: the middle value in a sorted version of the list

- Mean: arithmetic average of all values in the list

## Analog Conditioning Adjustment

The analog conditioning circuitry (±10V to ADC input voltage) shall provide the following adjustments for each channel (typically using a potentiometer):

- Zero: sets the offset voltage for the conversion (adjusts the "zero" voltage for the analog conditioning output).

- Span: sets the gain of the analog conditioning circuitry.

- VRef: sets the voltage reference for the single-ended to double-ended conversion. Note that the double-ended ADC inputs may only support voltages in the range 0-5V. Negative voltages are converted to a double-ended signal where the actual value is represented by two voltage inputs (both in the range 0-5V, Vp and Vn) where the resulting value is computed as (Vp-Vn). The VRef input determines the value where Vp=Vn to produce a zero result. This is different from the ZERO adjustment, which sets the input voltage offset.

## Host Communication

The System software shall communicate with the host computer using one of two different protocols: the DIS64 protocol or the PPDAQ protocol.

The DIS64 protocol is a legacy protocol compatible with the DIS64 (Digital Input Scanner, 64 bits) board found in older General Atomics digital console systems. This protocol provides a drop-in replacement for these obsolete boards in legacy systems. Note that the DIS64 protocol communicates over RS-232 serial at 9600 bits per second, no parity, and one stop bit (as per the DIS64 board). The DIS64 protocol is only implemented via the Teensy 3.2 computer module (to save space and development effort, which is important consider the small audience for this particular protocol). The DIS64 protocol is for legacy systems only, newer designs should not use this protocol.

The PPDAQ protocol is the "standard" communication protocol for the System. This protocol communicates with the host via the Ethernet, USB, or RS-232 Serial ports. This protocol is a human-readable ASCII-text protocol with each command consisting of a sequence of ASCII characters terminated by a new line sequence ("\n" – the line feed character). The exact form of the protocol appears in the SRS document.

## System Regulatory Issues

Although the DAQ System has been designed with modern hardware and software engineering principles in mind – with an eye towards achieving approval from appropriate governmental regulatory agencies – the installing site has the ultimate responsibility for ensuring that the DAQ System meets their needs and is in compliance with whatever regulations may apply.

Although each site is responsible for achieving approval for the use of the DAQ System, the hope is that the approval process will be reduced in scope as different sites can share their experiences getting the DAQ System approved at their site.

## User Characteristics

In general, there are three types of individuals who use the DAQ System: end users, technicans, and system engineers.

- System engineers are those who design the DAQ System into their particular reactor control system and are responsible for testing and verifying the proper operation of the DAQ System at their installation.

- Technicians are those who are responsible for maintaining and calibrating the DAQ System after it is installed at their site.

- End users are reactor operators who use the DAQ System (indirectly) whenever they operate the reactor. Note that reactor operators may be blissfully unaware that the DAQ System hardware even exists. The DAQ System hardware could be replaced (it it could replace older hardware) and the end-user may not even be aware of the change.

There is no requirement that these three types of people all be different. The system engineer could also be responsible for on-going maintenance and calibration and could also be a reactor operator (end user).

# System Requirements

## Physical Requirements

### Module Width

[DAQ_SYRS_001]

The circuit boards shall be 3.5" wide.

This will allow 3-4 boards across a drawer (in portrait mode) in a typical 19" rack-mount drawer.

Note: some connectors (e.g., RS-232) may hang over the edge of the board.

### Module Length

[DAQ_SYRS_002]

The circuit boards shall be 5.9" long.

This will allow 2-3 boards deep in a drawer (in portrait mode) in a typical 19" rack-mount drawer.

### Mounting Holes

[DAQ_SYRS_003]

Four mounting holes shall be centered 0.15" from each corner.

This implies the following:

- Holes are 3.20" apart (center to center) on the short side (width) of the board

- Holes are 5.6" apart (center to center) on the long side (length) of the board.

## Electromechanical Requirements

### Circuit Board Layers

[DAQ_SYRS_004]

The circuit boards shall be four layer boards.

### Circuit Board Top Layer

[DAQ_SYRS_005]

The circuit board top layer shall contain signal traces and auxiliary power planes.

### Circuit Board Second (Interior) Layer

[DAQ_SYRS_006]

The circuit board second layer shall contain a ground plane.

### Circuit Board Third (Interior) Layer

[DAQ_SYRS_007]

The circuit board third layer shall contain power (e.g., +5V,+12V, -12V) planes.

### Circuit Board Bottom Layer

[DAQ_SYRS_008]

The circuit board bottom layer shall contain signal traces and auxiliary power planes.

### Signal Trace Separation

[DAQ_SYRS_009]

Adjacent signal traces shall have a minimum of 8 mil separation.

### Minimum Trace Width

[DAQ_SYRS_010]

Signal traces shall have a minimum width of 8 mil.

### Minimum Power Trace Width

[DAQ_SYRS_011]

Power (and ground) traces shall have a minimum width of 15 mil.

### Minimum High-Current Trace Width

[DAQ_SYRS_012]

Traces carrying high-current signals (e.g., relay contacts) shall have a minimum width of 25 mil.

## Hardware Requirements

### CPU Interface Board (DAQIF)

#### Netburner MOD54415 DEV-70 Evaluation Board Support

[DAQ_SYRS_050]

The DAQIF board shall allow the connection of a Netburner MOD54415 (J2 Connector).

#### Raspberry Pi 3 Support

[DAQ_SYRS_051]

The DAQIF board shall allow the connection of a Raspberry Pi 3 (GPIO Connector).

#### Teensy 3.2 Support

[DAQ_SYRS_052]

The DAQIF board shall allow the connection of a Teensy 3.2 SBC (socket provided for headers on Teensy 3.2).

#### Single CPU Connection

[DAQ_SYRS_053]

The DAQIF board shall only support the connection of a single CPU board (Netburner, Pi, or Teensy).

#### 5V Logic Levels

[DAQ_SYRS_054]

The DAQIF board shall translate all voltage levels from the CPU boards to a 5V/TTL-compatible logic before such signals exit (or enter) the DAQIF .

Note: Some 3.3V output signals need not be converted as 3.3V outputs are TTL-logic compatible.

Note: I2C and SPI signals must be translated beween 3.3V and 5V before leaving the board.

#### TTL Fanout

[DAQ_SYRS_055]

All logic output signals from the DAQIF board shall support a TTL fanout of 10. (Buffer chips required if a particular signal has a fanout of less than 10.)

#### SPI Chip/Board Select Expansion

[DAQ_SYRS_056]

The DAQIF shall provide an SPI chip select expansion with 8 chip selects (BS0-BS7).

#### Reserved SPI Chip Select

[DAQ_SYRS_057]

The DAQIF shall reserve BS7 (chip select #8) for "nothing selected".

#### Serial Output SPI Chip Select

[DAQ_SYRS_058]

The DAQIF shall reserve BS0 (chip select #1) for a special serial output bus.

#### Parallel Input SPI Chip Selects

[DAQ_SYRS_059]

The DAQIF shall reserve BS1-BS6 (chip selects #2-#7) for accessing PPDIO96 parallel I/O boards.

#### I2C Expansion

[DAQ_SYRS_060]

The DAQIF shall provide an I2C expansion facility to expand a single I2C port to 8 I2C ports.

#### Watchdog Timer

[DAQ_SYRS_061]

The DAQIF shall provide watchdog timer to monitor software operation.

#### Watchdog Timeout Period

[DAQ_SYRS_062]

The DAQIF watchdog timeout period shall be between 2 and 10 seconds (nominally 5 seconds).

#### Watchdog Latch

[DAQ_SYRS_063]

The DAQIF shall latch the watchdog timeout output until the CPU explicitly resets it.

#### Watchdog Feedback

[DAQ_SYRS_064]

The DAQIF watchdog timeout signal shall be made available to the CPU module.

#### Watchdog Reset

[DAQ_SYRS_065]

The CPU module shall provide a digital output line that will reset the DAQIF watchdog trip.

Note: this reset line may be shared with other reset functions.

#### Watchdog Timeout Reset

[DAQ_SYRS_066]

The CPU module shall provide a digital output line that will reset the DAQIF watchdog timeout period.

#### Watchdog Relay

[DAQ_SYRS_067]

The DAQIF shall provide an SPDT relay with normally open and normally closed contacts.

#### Watchdog Relay – Not Actuated

[DAQ_SYRS_068]

The watchdog relay shall be unactuated (normally open contacts open, normally closed contacts closed) whenever the watchdog latch is not actuated.

#### Watchdog Relay – Actuated

[DAQ_SYRS_069]

The watchdog relay shall be actuated (normally open contacts closed, normally closed contacts open) whenever the watchdog latch is actuated.

#### Watchdog LED

[DAQ_SYRS_070]

The DAQIF shall provide an LED that illuminates whenever the watchdog latch is actuated.

#### Watchdog Output Connector

[DAQ_SYRS_071]

The DAQIF shall provide a 3-pin header (0.1" centers) with the NC/COM/NO signals from the watchdog relay.

#### Power Supply LED

[DAQ_SYRS_072]

The DAQIF shall provide an LED that is illuminated whenever power is applied to the board.

#### RS-232 Level Shifting

[DAQ_SYRS_073]

The DAQIF shall provide level shifting from TTL levels to RS-232 levels.

Note: this functionality is only necessary for the Raspberry Pi and Teensy CPUs. The Netburner MOD54415 evaluation board already provides this functionality.

#### RS-232 DCE Connection

[DAQ_SYRS_074]

The DAQIF shall provide a DB9F (female) connection with a DCE (data communications equipment) interface.

#### RS-232 DTE Connection

[DAQ_SYRS_075]

The DAQIF shall provide a DB9 (male) connection with a DTE (data terminal equipment) interface.

Note: DTE is the standard connector for this type of device. DCE is provided to avoid a NULL modem connection when connecting directly to a terminal device such as a PC.

#### PPDO Connection

[DAQ_SYRS_076]

The DAQIF shall provide a 10-pin header (2x5) to connect to serial output boards (e.g., PPSSR-16 and PPRELAY-12).

#### PPDO Connection Pinout

[DAQ_SYRS_077]

The DAQIF PPDO connector shall use the following pinout:

- BS0 (SPI chip select #1)

- VCC (+5V)

- MISO (5V SPI master data in, slave data out)

- MOSI (5V SPI master data out, slave data in)

- CLK (5V SPI clock)

- J2_38_5 (5V digital signal, reserved for future use)

- Reset (5V logic reset signal – common with watchdog reset, used to reset serial latches).

- Watchdog latch (5V signal indicates watchdog timeout).

- IRQ (5V—open drain signal to the CPU indicating an interrupt condition)

- Gnd

#### PPDIO96 Connection

[DAQ_SYRS_078]

The DAQIF shall provide a 14-pin header (2x7) to connect to PPDIO96 parallel I/O boards (or signal-compatible devices).

#### PPDIO96 Connection Pinout

[DAQ_SYRS_079]

The DAQIF  PPDIO96 connector shall use the following pinout:

- J2_38_5 (5V digital signal, reserved for future use)

- Watchdog latch (5V signal indicates watchdog timeout).

- BS1 (SPI chip select #2)

- Gnd

- BS2 (SPI chip select #3)

- MOSI (5V SPI master data out, slave data in)

- BS3 (SPI chip select #4)

- CLK (5V SPI Clock)

- BS4 (SPI chip select #5)

- MISO (5V SPI master data in, slave data out)

- BS5 (SPI chip select #6)

- IRQ (5V—open drain signal to the CPU indicating an interrupt condition)

- BS6 (SPI chip select #7)

- Reset (5V logic reset signal – common with watchdog reset, used to reset serial latches).

Note: BSn signals need to be TTL compatible; they could be 3.3V or 5V signals.

#### I2C Connections

[DAQ_SYRS_080]

The DAQIF shall four 6-pin (2x3) headers, each of which provides two I2C connections from the I2C expander.

#### I2C Connector Pinouts

[DAQ_SYRS_081]

The DAQIF  I2C connections shall use the following pinouts:

Connector 0/1:

- SD0 (5V I2C port 0 data)

- SD1 (5V I2C port 1 data)

- SC0 (5V I2C port 0 clock)

- SC1 (5V I2C port 1 clock)

- Gnd

- Gnd

Connector 2/3:

- SD2 (5V I2C port 2 data)

- SD3 (5V I2C port 3 data)

- SC2 (5V I2C port 2 clock)

- SC3 (5V I2C port 3 clock)

- Gnd

- Gnd

Connector 4/5:

- SD4 (5V I2C port 4 data)

- SD5 (5V I2C port 5 data)

- SC4 (5V I2C port 4 clock)

- SC5 (5V I2C port 5 clock)

- Gnd

- Gnd

Connector 6/7:

- SD6 (5V I2C port 6 data)

- SD7 (5V I2C port 7 data)

- SC6 (5V I2C port 6 clock)

- SC7 (5V I2C port 7 clock)

- Gnd

- Gnd

#### DAQIF Power Supply Connection

[DAQ_SYRS_082]

The DAQIF shall provide a two-pin screw terminal connector (5mm centers) to accept 5V/Gnd from an external power supply

#### DAQIF 3.3V Power Supply

[DAQ_SYRS_083]

The DAQIF shall provide on-board voltage regulation for any voltages other than the 5V provided on the power connection (e.g., 3.3V for many of the on-board logic chips).

### Parallel I/O Interface Board (PPDIO96)

#### PPDIO96 Capacity

[DAQ_SYRS_150]

The PPDIO96 shall provide 96 bits of parallel I/O.

#### PPDIO96 I/O Expansion Chips

[DAQ_SYRS_151]

The PPDIO96 shall use six MCP23S17 I/O expansion ICs (16 I/O pins per chip).

#### PPDIO96 Input Connection

[DAQ_SYRS_152]

The PPDIO96 shall use the PPDIO96 data connection (2x7 pin header, see [DAQ_SYRS_079]) as the input connector.

#### PPDIO96 Output Connection

[DAQ_SYRS_153]

The PPDIO96 shall use the PPDIO96 data connection (2x7 pin header, see [DAQ_SYRS_079]) as the output connector with the following changes:

- BS2 on input connector is routed to BS1 on output connector.

- BS3 on input connector is routed to BS2 on output connector.

- BS4 on input connector is routed to BS3 on output connector.

- BS5 on input connector is routed to BS4 on output connector.

- BS6 on input connector is routed to BS5 on output connector.

- BS6 on the output connector is tied to +5V.

#### PPDIO96 IRQ Connection

[DAQ_SYRS_154]

The PPDIO96 shall provide a jumper connection for each MCP23S17 chip that allows the user to connect the chip's INTA pin to the IRQ pin on the PPDIO96 connector.

#### PPDIO96 I/O Connections

[DAQ_SYRS_155]

The PPDIO96 shall provide eight 20-pin (2x10) headers connecting to the data pins on the MCP23S17 IC.

Note: this document refers to each of the connectors as a "bank".

#### PPDIO96 Bank 0 Pinout

[DAQ_SYRS_156]

The PPDIO96 Bank 0 header shall use the following pinout:

- Bit 1 (MCP23S17 #1 GPB6)

- Bit 0 (MCP23S17 #1 GPB7)

- Bit 2 (MCP23S17 #1 GPB5)

- Gnd

- Bit 3 (MCP23S17 #1 GPB4)

- Gnd

- Bit 4 (MCP23S17 #1 GPB3)

- Gnd

- Bit 5 (MCP23S17 #1 GPB2)

- Gnd

- Bit 6 (MCP23S17 #1 GPB1)

- Gnd

- Bit 7 (MCP23S17 #1 GPB0)

- Gnd

- Bit 8 (MCP23S17 #1 GPA7)

- Gnd

- Bit 9 (MCP23S17 #1 GPA6)

- Gnd

- Bit 10 (MCP23S17 #1 GPA5)

- Bit 11 (MCP23S17 #1 GPA4)

Note: Gnd pins are all connected to the same ground plane.

#### PPDIO96 Bank 1 Pinout

[DAQ_SYRS_157]

The PPDIO96 Bank 1 header shall use the following pinout:

- Bit 1 (MCP23S17 #1 GPA2)

- Bit 0 (MCP23S17 #1 GPA3)

- Bit 2 (MCP23S17 #1 GPA1)

- Gnd

- Bit 3 (MCP23S17 #1 GPA0)

- Gnd

- Bit 4 (MCP23S17 #2 GPB7)

- Gnd

- Bit 5 (MCP23S17 #2 GPB6)

- Gnd

- Bit 6 (MCP23S17 #2 GPB5)

- Gnd

- Bit 7 (MCP23S17 #2 GPB4)

- Gnd

- Bit 8 (MCP23S17 #2 GPB3)

- Gnd

- Bit 9 (MCP23S17 #2 GPB2)

- Gnd

- Bit 10 (MCP23S17 #2 GPB1)

- Bit 11 (MCP23S17 #2 GPB0)

Note: Gnd pins are all connected to the same ground plane.

#### PPDIO96 Bank 2 Pinout

[DAQ_SYRS_158]

The PPDIO96 Bank 2 header shall use the following pinout:

- Bit 1 (MCP23S17 #2 GPA1)

- Bit 0 (MCP23S17 #2 GPA0)

- Bit 2 (MCP23S17 #2 GPA2)

- Gnd

- Bit 3 (MCP23S17 #2 GPA3)

- Gnd

- Bit 4 (MCP23S17 #2 GPA4)

- Gnd

- Bit 5 (MCP23S17 #2 GPA5)

- Gnd

- Bit 6 (MCP23S17 #2 GPA6)

- Gnd

- Bit 7 (MCP23S17 #2 GPA7)

- Gnd

- Bit 8 (MCP23S17 #3 GPB7)

- Gnd

- Bit 9 (MCP23S17 #3 GPB6)

- Gnd

- Bit 10 (MCP23S17 #3 GPB5)

- Bit 11 (MCP23S17 #3 GPB4)

Note: Gnd pins are all connected to the same ground plane.

#### PPDIO96 Bank 3 Pinout

[DAQ_SYRS_159]

The PPDIO96 Bank 3 header shall use the following pinout:

- Bit 1 (MCP23S17 #3 GPB2)

- Bit 0 (MCP23S17 #3 GPB3)

- Bit 2 (MCP23S17 #3 GPB1)

- Gnd

- Bit 3 (MCP23S17 #3 GPB0)

- Gnd

- Bit 4 (MCP23S17 #3 GPA7)

- Gnd

- Bit 5 (MCP23S17 #3 GPA6)

- Gnd

- Bit 6 (MCP23S17 #3 GPA5)

- Gnd

- Bit 7 (MCP23S17 #3 GPA4)

- Gnd

- Bit 8 (MCP23S17 #3 GPA3)

- Gnd

- Bit 9 (MCP23S17 #3 GPA2)

- Gnd

- Bit 10 (MCP23S17 #3 GPA1)

- Bit 11 (MCP23S17 #3 GPA0)

Note: Gnd pins are all connected to the same ground plane.

#### PPDIO96 Bank 4 Pinout

[DAQ_SYRS_160]

The PPDIO96 Bank 4 header shall use the following pinout:

- Bit 1 (MCP23S17 #4 GPB2)

- Bit 0 (MCP23S17 #4 GPB3)

- Bit 2 (MCP23S17 #4 GPB1)

- Gnd

- Bit 3 (MCP23S17 #4 GPB0)

- Gnd

- Bit 4 (MCP23S17 #4 GPA7)

- Gnd

- Bit 5 (MCP23S17 #4 GPA6)

- Gnd

- Bit 6 (MCP23S17 #4 GPA5)

- Gnd

- Bit 7 (MCP23S17 #4 GPA4)

- Gnd

- Bit 8 (MCP23S17 #4 GPA3)

- Gnd

- Bit 9 (MCP23S17 #4 GPA2)

- Gnd

- Bit 10 (MCP23S17 #4 GPA1)

- Bit 11 (MCP23S17 #4 GPA0)

Note: Gnd pins are all connected to the same ground plane.

#### PPDIO96 Bank 5 Pinout

[DAQ_SYRS_161]

The PPDIO96 Bank 5 header shall use the following pinout:

- Bit 1 (MCP23S17 #5 GPA6)

- Bit 0 (MCP23S17 #5 GPA7)

- Bit 2 (MCP23S17 #5 GPA5)

- Gnd

- Bit 3 (MCP23S17 #5 GPA4)

- Gnd

- Bit 4 (MCP23S17 #5 GPA3)

- Gnd

- Bit 5 (MCP23S17 #5 GPA2)

- Gnd

- Bit 6 (MCP23S17 #5 GPA1)

- Gnd

- Bit 7 (MCP23S17 #5 GPA0)

- Gnd

- Bit 8 (MCP23S17 #4 GPB7)

- Gnd

- Bit 9 (MCP23S17 #4 GPB6)

- Gnd

- Bit 10 (MCP23S17 #4 GPB5)

- Bit 11 (MCP23S17 #4 GPB4)

Note: Gnd pins are all connected to the same ground plane.

#### PPDIO96 Bank 6 Pinout

[DAQ_SYRS_162]

The PPDIO96 Bank 6 header shall use the following pinout:

- Bit 1 (MCP23S17 #6 GPA2)

- Bit 0 (MCP23S17 #6 GPA3)

- Bit 2 (MCP23S17 #6 GPA1)

- Gnd

- Bit 3 (MCP23S17 #6 GPA0)

- Gnd

- Bit 4 (MCP23S17 #5 GPB7)

- Gnd

- Bit 5 (MCP23S17 #5 GPB6)

- Gnd

- Bit 6 (MCP23S17 #5 GPB5)

- Gnd

- Bit 7 (MCP23S17 #5 GPB4)

- Gnd

- Bit 8 (MCP23S17 #5 GPB3)

- Gnd

- Bit 9 (MCP23S17 #5 GPB2)

- Gnd

- Bit 10 (MCP23S17 #5 GPB1)

- Bit 11 (MCP23S17 #5 GPB0)

Note: Gnd pins are all connected to the same ground plane.

#### PPDIO96 Bank 7 Pinout

[DAQ_SYRS_163]

The PPDIO96 Bank 7 header shall use the following pinout:

- Bit 1 (MCP23S17 #6 GPB6)

- Bit 0 (MCP23S17 #6 GPB7)

- Bit 2 (MCP23S17 #6 GPB5)

- Gnd

- Bit 3 (MCP23S17 #6 GPB4)

- Gnd

- Bit 4 (MCP23S17 #6 GPB3)

- Gnd

- Bit 5 (MCP23S17 #6 GPB2)

- Gnd

- Bit 6 (MCP23S17 #6 GPB1)

- Gnd

- Bit 7 (MCP23S17 #6 GPB0)

- Gnd

- Bit 8 (MCP23S17 #6 GPA7)

- Gnd

- Bit 9 (MCP23S17 #6 GPA6)

- Gnd

- Bit 10 (MCP23S17 #6 GPA5)

- Bit 11 (MCP23S17 #6 GPA4)

Note: Gnd pins are all connected to the same ground plane.

#### PPDIO96 DIP Switches

[DAQ_SYRS_164]

The PPDIO96 shall provide 8 DIP switches connected to bits 88-95 (bank 7, bits 4-11).

#### PPDIO96 Power Supply Connection

[DAQ_SYRS_165]

The PPDIO96 shall provide a two-pin screw terminal connector (5mm centers) to accept 5V/Gnd from an external power supply

#### PPDIO96 Power Supply LED

[DAQ_SYRS_166]

The PPDIO96 shall provide an LED that is illuminated whenever power is applied to the board.

#### PPDIO96 Buffered Inputs

[DAQ_SYRS_167]

The PPDIO96 shall provide buffering so that it presents a single TTL load to the following signals:

- MOSI

- SPI-CLK

#### PPDIO96 Buffered Output

[DAQ_SYRS_168]

The PPDIO96 shall provide buffering for the MISO output line

#### PPDIO96 Tri-state Buffering

[DAQ_SYRS_169]

The PPDIO96 buffering shall be tri-state with the buffer output lines active only when the BS1 signal is active.

### 12-Channel Optoisolator Parallel Input Board (PPOPTO-12)

#### PPOPTO-12 Capacity

[DAQ_SYRS_200]

The PPOPTO-12 shall provide 12 bits of optically-isolated digital inputs.

#### PPOPTO-12 Power Supply Connection

[DAQ_SYRS_201]

The PPOPTO-12  board shall provide a two-pin screw terminal connector (5mm centers) to accept 5V/Gnd from an external power supply

#### PPOPTO-12 Power Supply LED

[DAQ_SYRS_202]

The PPOPTO-12  board shall provide an LED that is illuminated whenever power is applied to the board.

#### PPOPTO-12 Output Connector Pinout

[DAQ_SYRS_203]

The PPOPTO-12 board shall provide a 20-pin (2x10) header that is compatible with the PPDIO96 bank pinouts, using the following pinout:

- Bit 1

- Bit 0

- Bit 2

- Gnd

- Bit 3

- Gnd

- Bit 4

- Gnd

- Bit 5

- Gnd

- Bit 6

- Gnd

- Bit 7

- Gnd

- Bit 8

- Gnd

- Bit 9

- Gnd

- Bit 10

- Bit 11

Note: Gnd pins are all connected to the same ground plane.

#### PPOPTO-12 Signal LEDs

[DAQ_SYRS_204]

The PPOPTO-12 board shall provide 12 signal LEDs.

#### PPOPTO-12 Signal LED #0 Illumination

[DAQ_SYRS_205]

Signal LED #0 on the PPOPTO-12 board shall illuminate when input D0 is active.

#### PPOPTO-12 Signal LED #1 Illumination

[DAQ_SYRS_206]

Signal LED #1 on the PPOPTO-12 board shall illuminate when input D1 is active.

#### PPOPTO-12 Signal LED #2 Illumination

[DAQ_SYRS_207]

Signal LED #2 on the PPOPTO-12 board shall illuminate when input D2 is active.

#### PPOPTO-12 Signal LED #3 Illumination

[DAQ_SYRS_208]

Signal LED #3 on the PPOPTO-12 board shall illuminate when input D3 is active.

#### PPOPTO-12 Signal LED #4 Illumination

[DAQ_SYRS_209]

Signal LED #4 on the PPOPTO-12 board shall illuminate when input D4 is active.

#### PPOPTO-12 Signal LED #5 Illumination

[DAQ_SYRS_210]

Signal LED #5 on the PPOPTO-12 board shall illuminate when input D5 is active.

#### PPOPTO-12 Signal LED #6 Illumination

[DAQ_SYRS_211]

Signal LED #6 on the PPOPTO-12 board shall illuminate when input D6 is active.

#### PPOPTO-12 Signal LED #7 Illumination

[DAQ_SYRS_212]

Signal LED #7 on the PPOPTO-12 board shall illuminate when input D7 is active.

#### PPOPTO-12 Signal LED #8 Illumination

[DAQ_SYRS_213]

Signal LED #8 on the PPOPTO-12 board shall illuminate when input D8 is active.

#### PPOPTO-12 Signal LED #9 Illumination

[DAQ_SYRS_214]

Signal LED #9 on the PPOPTO-12 board shall illuminate when input D9 is active.

#### PPOPTO-12 Signal LED #10 Illumination

[DAQ_SYRS_215]

Signal LED #10 on the PPOPTO-12 board shall illuminate when input D10 is active.

#### PPOPTO-12 Signal LED #11 Illumination

[DAQ_SYRS_216]

Signal LED #11 on the PPOPTO-12 board shall illuminate when input D11 is active.

#### PPOPTO-12 Electrical Isolation

[DAQ_SYRS_217]

Each input channel on the PPOPTO-12 board shall electrically isolated from the other input channels.

Note: this implies that each channel must have an isolated power supply in order to work with dry contact inputs.

#### PPOPTO-12 Dry Contact Input Selection

[DAQ_SYRS_218]

Each input channel on the PPOPTO-12 board shall offer the ability to work with a dry contact input.

Note: Dry contact inputs require an isolated power supply for each input channel.

#### PPOPTO-12 Voltage Input Selection

[DAQ_SYRS_219]

Each input channel on the PPOPTO-12 board shall offer the ability to accept a voltage logic-level input (supplied by the external circuitry).

#### PPOPTO-12 Voltage Input Range

[DAQ_SYRS_220]

Each input channel on the PPOPTO-12 board shall be able to work with voltage inputs in the range 0-24V.

Note: programming the input voltage may necessitate changing a resistor.

#### PPOPTO-12 Input Connectors

[DAQ_SYRS_221]

The PPOPTO-12 board shall provide a two-pin screw terminal for each input (5mm centers).

### 16-Channel Solid-State Relay Output Board (PPSSR16)

#### PPSSR16 Capacity

[DAQ_SYRS_250]

The PPSSR16 board shall provide 16 output channels (bits).

#### PPSSR16 Power Supply Connection

[DAQ_SYRS_251]

The PPSSR16 board shall provide a two-pin screw terminal connector (5mm centers) to accept 5V/Gnd from an external power supply

#### PPSSR16 Power Supply LED

[DAQ_SYRS_252]

The PPSSR16 board shall provide an LED that is illuminated whenever power is applied to the board.

#### PPSSR16 Input Interface Connection

[DAQ_SYRS_253]

The PPSSR16 board shall communicate with the DAQIF using the PPDO connector on the DAQIF (replicated as input connector on the PPSSR16 board); the pinout is described in [DAQ_SYRS_077]

#### PPSSR16 Daisy Chain Connection

[DAQ_SYRS_254]

The PPSSR16 board shall pass through the following signals from the input connector to the output connector:

- BSO (SPI bus chip select #0)

- NC

- MISO

- – (not passed through, see [DAQ_SYRS_255], [DAQ_SYRS_258], and [DAQ_SYRS_259])

- SPI-CLK

- J2_38

- J2_41 (Watchdog timeout)

- IRQ

- Gnd

#### PPSSR16 Daisy Chaining

[DAQ_SYRS_255]

The PPSSR16 board shall pass the SER OUT output of the second on-board TPIC6B595PTH  serial shift register through to the MOSI pin (pin 4) on the output PPDO connection.

#### PPSSR16 BS0 Signal

[DAQ_SYRS_256]

The PPSSR16 board shall use the BS0 (SPI bus board/chip select #0) signal on the PPDO connector to control the transfer of the data from the internal shift register to the output latch on the TPIC6B595PTH high-current serial shift register.

Note: data is transferred from the shift register to the output latch on the rising edge of this signal – which occurs when the system completes shifting all the data out to the shift register.

#### PPSSR16 MISO Signal

[DAQ_SYRS_257]

The PPSSR16 board shall ignore the MISO (master data in, slave out) signal on the PPDO connector. (Other than to pass it through to the PPDO output connector.)

#### PPSSR16 MOSI Signal

[DAQ_SYRS_258]

The PPSSR16 board shall connect the MOSI (SPI master out, slave in) input signal to the serial in pin of the first TPIC6B595PTH serial shift register.

#### PPSSR16 TPIC6B595PTH Daisy-Chaining

[DAQ_SYRS_259]

The PPSSR16 board shall connect the SER OUT signal of the first TPIC6B595PTH IC to the SER IN signal of the second TPIC6B595PTH IC.

#### PPSSR16 TPIC6B595PTH Power-on Reset

[DAQ_SYRS_260]

The PPSSR16 board shall clear the TPIC6B595PTH serial shift register on power up.

#### PPSSR16 TPIC6B595PTH Programmed Reset

[DAQ_SYRS_261]

The PPSSR16 board shall clear the TPIC6B595PTH serial shift registers whenever the reset signal on the PPDO connection is asserted (active low).

Note: The reset signal on the PPDO connector is shared amongst many functions, including watchdog latch reset and MCP23S17 (on the PPDIO96) reset.

#### PPSSR16 TPIC6B595PTH Watchdog Reset

[DAQ_SYRS_262]

The PPSSR16 board shall clear the TPIC6B595PTH serial shift registers whenever the WS Latch signal on the PPDO connection is asserted (active low).

#### PPSSR16 Data Clock

[DAQ_SYRS_263]

The PPSSR16 board shall connect the SPI-CLK signal on the PPDO connector to the SRCK (shift register clock) on both the TPIC6B595PTH  ICs.

Note: because the two TPIC6B595PTH  ICs are connected in serial (via the SER OUT and SER In pins), it will take 16 clock pulses to shift in 16 bits of data.

#### PPSSR16 Data Output LEDs

[DAQ_SYRS_264]

The PPSSR16 board shall provide one LED for each output channel.

#### PPSSR16 Data Output LED Illumination

[DAQ_SYRS_265]

The PPSSR16 board shall turn on a channel's output LED when that output channel is programmed with a '1' bit.

#### PPSSR16 Output Isolation

[DAQ_SYRS_266]

The PPSSR16 board shall provide isolation for all outputs.

#### PPSSR16 Output Current Capacity

[DAQ_SYRS_267]

The PPSSR16 board shall have relays capable of handling 2A.

#### PPSSR16 Output Voltage Capacity

[DAQ_SYRS_268]

The PPSSR16 board shall be capable of handling 30VDC on each output channel.

#### PPSSR16 Output Connectors

[DAQ_SYRS_269]

The PPSSR16 board shall provide a two-pin screw terminal for each relay output (5mm centers).

#### PPSSR16 Relay Sockets

[DAQ_SYRS_270]

The PPSSR16 board shall allow sockets for each solid-state relay on the board.

(Deprecated)

#### PPSSR16 Buffered SPI Clock signal

[DAQ_SYRS_271]

The PPSSR16 board shall buffer the incoming SPI-CLK (SCK) signal so that the board presents only one TTL load for this signal.

#### PPSSR16 Buffered RESET signal

[DAQ_SYRS_272]

The PPSSR16 board shall buffer the incoming RESET signal so that the board presents only one TTL load for this signal.

#### PPSSR16 Buffered Watchdog Latch signal

[DAQ_SYRS_273]

The PPSSR16 board shall buffer the incoming WD_Latch signal so that the board presents only one TTL load for this signal.

#### PPSSR16 Buffered Board Select signal

[DAQ_SYRS_274]

The PPSSR16 board shall buffer the incoming BS0 signal so that the board presents only one TTL load for this signal.

### 12-Channel Mechanical Relay Output Board (PPRELAY12)

#### PPRELAY12 Capacity

[DAQ_SYRS_300]

The PPRELAY12  board shall provide at least 12 output channels (bits).

#### PPRELAY12 Power Supply Connection

[DAQ_SYRS_301]

The PPRELAY12 board shall provide a two-pin screw terminal connector (5mm centers) to accept 5V/Gnd from an external power supply

#### PPRELAY12 Power Supply LED

[DAQ_SYRS_302]

The PPRELAY12 board shall provide an LED that is illuminated whenever power is applied to the board.

#### PPRELAY12 Input Interface Connection

[DAQ_SYRS_303]

The PPRELAY12 board shall communicate with the DAQIF using the PPDO connector on the DAQIF (replicated as input connector on the PPRELAY12 board); the pinout is described in [DAQ_SYRS_077]

#### PPRELAY12 Daisy Chain Connection

[DAQ_SYRS_304]

The PPRELAY12 board shall pass through the following signals from the input connector to the output connector:

- BSO (SPI bus chip select #0)

- NC

- MISO

- – (not passed through, see [DAQ_SYRS_305], [DAQ_SYRS_308], and [DAQ_SYRS_309])

- SPI-CLK

- J2_38

- Reset

- WD (Watchdog timeout)

- IRQ

- Gnd

#### PPRELAY12 Daisy Chaining

[DAQ_SYRS_305]

The PPRELAY12 board shall pass the SER OUT output of the second on-board TPIC6B595PTH  serial shift register through to the MOSI pin (pin 4) on the output PPDO connection.

#### PPRELAY12 BS0 Signal

[DAQ_SYRS_306]

The PPRELAY12 board shall use the BS0 (SPI bus board/chip select #0) signal on the PPDO connector to control the transfer of the data from the internal shift register to the output latch on the TPIC6B595PTH high-current serial shift register.

Note: data is transferred from the shift register to the output latch on the rising edge of this signal – which occurs when the system completes shifting all the data out to the shift register.

#### PPRELAY12 MISO Signal

[DAQ_SYRS_307]

The PPRELAY12 board shall ignore the MISO (master data in, slave out) signal on the PPDO connector.

#### PPRELAY12 MOSI Signal

[DAQ_SYRS_308]

The PPRELAY12 board shall connect the MOSI (SPI master out, slave in) input signal to the serial in pin of the first TPIC6B595PTH serial shift register.

#### PPRELAY12 TPIC6B595PTH Daisy-Chaining

[DAQ_SYRS_309]

The PPRELAY12 board shall connect the SER OUT signal of the first TPIC6B595PTH IC to the SER IN signal of the second TPIC6B595PTH IC.

#### PPRELAY12 TPIC6B595PTH Power-on Reset

[DAQ_SYRS_310]

The PPRELAY12 board shall clear the TPIC6B595PTH serial shift register on power up.

#### PPRELAY12 TPIC6B595PTH Programmed Reset

[DAQ_SYRS_311]

The PPRELAY12 board shall clear the TPIC6B595PTH serial shift register whenever the reset signal on the PPDO connection is asserted (active low).

Note: The reset signal on the PPDO connector is shared amongst many functions, including watchdog latch reset and MCP23S17 (on the PPDIO96) reset.

#### PPRELAY12 Watchdog Latch Reset

[DAQ_SYRS_312]

The PPRELAY12 board shall clear the TPIC6B595PTH serial shift register whenever the WD Latch signal on the PPDO connection is asserted (active low).

#### PPRELAY12 Data Clock

[DAQ_SYRS_313]

The PPRELAY12 board shall connect the SPI-CLK signal on the PPDO connector to the SRCK (shift register clock) on both the TPIC6B595PTH  ICs.

Note: because the two TPIC6B595PTH  ICs are connected in serials (via the SER OUT and SER In pins), it will take 16 clock pulses to shift in 16 bits of data.

#### PPRELAY12 Data Output LEDs

[DAQ_SYRS_314]

The PPRELAY12 board shall provide one LED for each output channel.

#### PPRELAY12 Data Output LED Illumination

[DAQ_SYRS_315]

The PPRELAY12 board shall turn on a channel's output LED when that output channel is programmed with a '1' bit.

#### PPRELAY12 Auxiliary Output Bits Connector

[DAQ_SYRS_316]

The PPRELAY12 board shall route the extra 4 bits (TPIC6B595PTH  ICs produce 16 output bits, there are only 12 relays) to an 8-pin connector (2x4) with the following pinout:

- Bit 12 output

- Vcc (5V)

- Bit 13 output

- Vcc (5V)

- Bit 14 output

- Vcc (5V)

- Bit 15 output

- Vcc (5V)

Note: the TPIC6B595PTH  IC provides open-drain outputs and pulls the signal to ground when active. This is why the alternate pins have +5V applied rather than ground (the TPIC6B595PTH  pins serve as the ground).

#### PPRELAY12 Output Isolation

[DAQ_SYRS_317]

The PPSSR16 board shall provide isolation for all outputs.

Note: this feature is provided by default given the nature of mechanical relays.

#### PPRELAY12 Output Current Capacity

[DAQ_SYRS_318]

The PPRELAY12 board shall be capable of handling an minimum of 2A on each output channel.

#### PPRELAY12 Output Voltage Capacity

[DAQ_SYRS_319]

The PPRELAY12 board shall be capable of handling a minimum of  30VDC on each output channel.

#### PPRELAY12 Output Connectors

[DAQ_SYRS_320]

The PPRELAY12 board shall provide a three-pin screw terminal for each relay output (5mm centers) for NC/COMM/NO relay connections.

#### PPRELAY12 NC Optional

[DAQ_SYRS_321]

The PPRELAY12 board shall provide the option of using SPST relays (with NO/COM connections only) by swapping relays and substituting two-terminal connections in place of the three-terminal connections ([DAQ_SYRS_319]).

#### PPRELAY12 Buffered SPI Clock signal

[DAQ_SYRS_322]

The PPRELAY12 board shall buffer the incoming SPI-CLK (SCK) signal so that the board presents only one TTL load for this signal.

#### PPRELAY12 Buffered RESET signal

[DAQ_SYRS_323]

The PPRELAY12 board shall buffer the incoming RESET signal so that the board presents only one TTL load for this signal.

#### PPRELAY12 Buffered Watchdog Latch signal

[DAQ_SYRS_324]

The PPRELAY12 board shall buffer the incoming WD_Latch signal so that the board presents only one TTL load for this signal.

#### PPRELAY12 Buffered BS0 signal

[DAQ_SYRS_325]

The PPRELAY12 board shall buffer the incoming BS0 signal so that the board presents only one TTL load for this signal.

### 48-Channel Digital Output Board (PPDO-48)

#### PPDO-48 Capacity

[DAQ_SYRS_330]

The PPDO-48 shall provide 48 bits of digital output.

#### PPDO-48 Power Supply Connection

[DAQ_SYRS_331]

The PPDO-48 board shall provide a two-pin screw terminal connector (5mm centers) to accept 5V/Gnd from an external power supply

#### PPDO-48 Power Supply LED

[DAQ_SYRS_332]

The PPDO-48 board shall provide an LED that is illuminated whenever power is applied to the board.

#### PPDO-48 Input Interface Connection

[DAQ_SYRS_333]

The PPDO-48 board shall communicate with the DAQIF using the PPDO connector on the DAQIF (replicated as input connector on the PPDO-48 board); the pinout is described in [DAQ_SYRS_077]

#### PPDO-48 Daisy Chain Connection

[DAQ_SYRS_334]

The PPDO-48 board shall pass through the following signals from the input connector to the output connector:

- BSO (SPI bus chip select #0)

- NC

- MISO

- – (not passed through, see [DAQ_SYRS_305], [DAQ_SYRS_308], and [DAQ_SYRS_309])

- SPI-CLK

- J2_38

- Reset

- WD (Watchdog timeout)

- IRQ

- Gnd

#### PPDO-48 Daisy Chaining

[DAQ_SYRS_335]

The PPDO-48 board shall pass the SER OUT output of the last on-board TPIC6B595PTH  serial shift register through to the MOSI pin (pin 4) on the output PPDO connection.

#### PPDO-48 BS0 Signal

[DAQ_SYRS_336]

The PPDO-48 board shall use the BS0 (SPI bus board/chip select #0) signal on the PPDO connector to control the transfer of the data from the internal shift register to the output latch on the TPIC6B595PTH high-current serial shift register.

Note: data is transferred from the shift register to the output latch on the rising edge of this signal – which occurs when the system completes shifting all the data out to the shift register.

#### PPDO-48 MISO Signal

[DAQ_SYRS_337]

The PPDO-48 board shall ignore the MISO (master data in, slave out) signal on the PPDO connector.

#### PPDO-48 MOSI Signal

[DAQ_SYRS_338]

The PPDO-48 board shall connect the MOSI (SPI master out, slave in) input signal to the serial in pin of the first TPIC6B595PTH serial shift register.

#### PPDO-48 TPIC6B595PTH Daisy-Chaining

[DAQ_SYRS_339]

The PPDO-48 board shall connect the SER OUT signal of the first TPIC6B595PTH IC to the SER IN signal of the second TPIC6B595PTH IC.

[DAQ_SYRS_339.1]

The PPDO-48 board shall connect the SER OUT signal of the second TPIC6B595PTH IC to the SER IN signal of the third TPIC6B595PTH IC.

[DAQ_SYRS_339.2]

The PPDO-48 board shall connect the SER OUT signal of the third TPIC6B595PTH IC to the SER IN signal of the fourth TPIC6B595PTH IC.

[DAQ_SYRS_339.3]

The PPDO-48 board shall connect the SER OUT signal of the fourth TPIC6B595PTH IC to the SER IN signal of the fifth TPIC6B595PTH IC.

[DAQ_SYRS_339.4]

The PPDO-48 board shall connect the SER OUT signal of the fifth TPIC6B595PTH IC to the SER IN signal of the sixth (last) TPIC6B595PTH IC.

#### PPDO-48 TPIC6B595PTH Power-on Reset

[DAQ_SYRS_340]

The PPDO-48 board shall clear the TPIC6B595PTH serial shift registers on power up.

#### PPRELAY12 TPIC6B595PTH Programmed Reset

[DAQ_SYRS_341]

The PPDO-48 board shall clear the TPIC6B595PTH serial shift registers whenever the reset signal on the PPDO connection is asserted (active low).

Note: The reset signal on the PPDO connector is shared amongst many functions, including watchdog latch reset and MCP23S17 (on the PPDIO96) reset.

#### PPDO-48 Watchdog Latch Reset

[DAQ_SYRS_342]

The PPDO-48 board shall clear the TPIC6B595PTH serial shift register whenever the WD Latch signal on the PPDO connection is asserted (active low).

#### PPDO-48 Data Clock

[DAQ_SYRS_343]

The PPDO-48 board shall connect the SPI-CLK signal on the PPDO connector to the SRCK (shift register clock) on all the TPIC6B595PTH ICs.

Note: because the six TPIC6B595 ICs are connected in series (via the SER OUT and SER In pins), it will take 48 clock pulses to shift in 48 bits of data.

#### PPDO-48 Output Connectors

[DAQ_SYRS_344.001]

The PPDO-48 board shall support six 16-pin output connectors (one for each TPIC6B595 IC on the board).

#### PPDO-48 Output Connector (Bank 0)

[DAQ_SYRS_344.002]

The first data output connector on the PPDO-48 shall support the following pinout:

- Vcc

- D0

- Vcc

- D1

- Vcc

- D2

- Vcc

- D3

- Gnd

- D4

- Gnd

- D5

- Gnd

- D6

- Gnd

- D7

#### PPDO-48 Output Connector (Bank 1)

[DAQ_SYRS_344.003]

The second data output connector on the PPDO-48 shall support the following pinout:

- Vcc

- D8

- Vcc

- D9

- Vcc

- D10

- Vcc

- D11

- Gnd

- D12

- Gnd

- D13

- Gnd

- D14

- Gnd

- D15

#### PPDO-48 Output Connector (Bank 2)

[DAQ_SYRS_344.004]

The third data output connector on the PPDO-48 shall support the following pinout:

- Vcc

- D16

- Vcc

- D17

- Vcc

- D18

- Vcc

- D19

- Gnd

- D20

- Gnd

- D21

- Gnd

- D22

- Gnd

- D23

#### PPDO-48 Output Connector (Bank 3)

[DAQ_SYRS_344.005]

The fourth data output connector on the PPDO-48 shall support the following pinout:

- Vcc

- D24

- Vcc

- D25

- Vcc

- D26

- Vcc

- D27

- Gnd

- D28

- Gnd

- D29

- Gnd

- D30

- Gnd

- D31

#### PPDO-48 Output Connector (Bank 4)

[DAQ_SYRS_344.006]

The fifth data output connector on the PPDO-48 shall support the following pinout:

- Vcc

- D32

- Vcc

- D33

- Vcc

- D34

- Vcc

- D35

- Gnd

- D36

- Gnd

- D37

- Gnd

- D38

- Gnd

- D39

#### PPDO-48 Output Connector (Bank 5)

[DAQ_SYRS_344.007]

The sixth data output connector on the PPDO-48 shall support the following pinout:

- Vcc

- D40

- Vcc

- D41

- Vcc

- D42

- Vcc

- D43

- Gnd

- D44

- Gnd

- D45

- Gnd

- D46

- Gnd

- D47

#### PPDO-48 Buffered SPI Clock signal

[DAQ_SYRS_344.008]

The PPDO-48 board shall buffer the incoming SPI-CLK (SCK) signal so that the board presents only one TTL load for this signal.

#### PPDO-48 Buffered RESET signal

[DAQ_SYRS_344.009]

The PPDO-48 board shall buffer the incoming RESET signal so that the board presents only one TTL load for this signal.

#### PPDO-48 Buffered Watchdog Latch signal

[DAQ_SYRS_344.010]

The PPDO-48 board shall buffer the incoming WD_Latch signal so that the board presents only one TTL load for this signal.

#### PPDO-48 Buffered BS0 signal

[DAQ_SYRS_344.011]

The PPDO-48 board shall buffer the incoming BS0 signal so that the board presents only one TTL load for this signal.

### 16-Channel Analog Input/4-Channel Analog Output Board (PPAIO-16/4)

#### PPAIO-16/4 Analog Input Capacity

[DAQ_SYRS_350]

The PPAIO-16/4 board shall provide 16 singled-ended analog input channels.

#### PPAIO-16/4 Double-Ended Input Capacity

[DAQ_SYRS_351]

The PPAIO-16/4 board shall allow adjacent even-odd pairs of single-ended analog inputs to be combined into a double-ended input.

Note: this implies a maximum of eight double-ended inputs.

#### PPAIO-16/4 Analog Output Capacity

[DAQ_SYRS_352]

The PPAIO-16/4 board shall provide 4 single-ended analog output channels.

#### PPAIO-16/4 ADC Input Voltage Range

[DAQ_SYRS_353]

The PPAIO-16/4 board shall accept input voltages in the range 0-4.096V on the ADC input pins.

Note: this applies to both single-ended inputs and double-ended inputs

#### PPAIO-16/4 Analog Input Resolution (double-ended)

[DAQ_SYRS_354]

The PPAIO-16/4 board shall support 16-bit 2's complement resolution on the analog inputs when using double-ended inputs.

Note: The ADC shall return negative values when the voltage on the double-ended negative input pin is greater than the voltage on the double-ended positive input pin.

#### PPAIO-16/4 Analog Input Resolution (single-ended)

[DAQ_SYRS_355]

The PPAIO-16/4 board shall support 15-bit  analog inputs when using single-ended inputs.

Note: value is 0-0x7FFF (0-32767) based on a voltage input of 0-4.096 volts on the single-ended input.

#### PPAIO-16/4 ADC Conversion Rate

[DAQ_SYRS_356]

The PPAIO-16/4 board shall be capable of producing a minimum of 100 samples/sec.

#### PPAIO-16/4 Analog Output Capacity

[DAQ_SYRS_357]

The PPAIO-16/4 board shall provide 4 singled-ended analog output channels.

#### PPAIO-16/4 DAC Output Voltage Range

[DAQ_SYRS_358]

The PPAIO-16/4 board shall produce output voltages in the range -5V to +5V on the DAC output pins.

#### PPAIO-16/4 Analog Output Resolution

[DAQ_SYRS_359]

The PPAIO-16/4 board shall support 12-bit  analog outputs. Values in the range 0-4095 produce a voltage in the range -5V to 5V (2048 corresponds to 0V).

#### PPAIO-16/4 DAC Conversion Rate

[DAQ_SYRS_360]

The PPAIO-16/4 board shall be capable of producing a minimum of 100 output samples/sec.

#### PPAIO-16/4 Digital Input Connections

[DAQ_SYRS_361]

The PPAIO-16/4 board shall provide a 6-pin (2x3) header accepting two sets of I2C control signals. The pinout is:

SDA (5V I2C port A data)

SDB (5V I2C port B data)

SCA (5V I2C port A clock)

SCB (5V I2C port B clock)

Gnd

Gnd

#### PPAIO-16/4 Digital Output Connections

[DAQ_SYRS_362]

The PPAIO-16/4 board shall provide a 6-pin (2x3) header used to daisy-chain a second  PPAIO-16/4 board. The pin configuration of the header is similar to the digital input connection except that it swaps the pins for the two i2c busses. The pin out is:

- SDB (5V I2C port B data)

- SDA (5V I2C port A data)

- SCB (5V I2C port B clock)

- SCA (5V I2C port A clock)

- Gnd

- Gnd

Note: a maximum of two PPAIO-16/4 boards can be daisy-chained together using this scheme. Attempting to daisy-chain a third board reverts back to the A bus and creates address conflicts between the first and third boards.

#### PPAIO-16/4 Analog Input Connections

[DAQ_SYRS_363]

The PPAIO-16/4 board shall provide 16 two-terminal screw terminals corresponding to inputs AIN0-AIN15 (analog inputs 0-15).

#### PPAIO-16/4 Analog Output Connections

[DAQ_SYRS_364]

The PPAIO-16/4 board shall provide 4 two-terminal screw terminals corresponding to outputs AOUT0-AOUT3 (analog outputs 0-3).

#### PPAIO-16/4 ADC Modules

[DAQ_SYRS_365]

The PPAIO-16/4 board shall support four Adafruit ADS1115 16-bit ADC modules.

#### PPAIO-16/4 Adafruit ADS1115 I2C Usage

[DAQ_SYRS_366]

The four Adafruit ADS1115 16-bit ADC modules shall only use the first I2C signals on the digital input connection (pins 1 and 3 on the connector, SDA and SCA).

#### PPAIO-16/4 ADC Module #0 Address Selection

[DAQ_SYRS_367]

The first Adafruit ADS1115 module shall be assigned address zero.

Note: Address zero is obtained by tying the ADDR pin to Gnd.

#### PPAIO-16/4 ADC Module #1 Connection to AIN0

[DAQ_SYRS_368]

The first Adafruit ADS1115 module A0 input shall be connected to the positive side of the AIN0 screw terminal.

#### PPAIO-16/4 ADC Module #1 Connection to AIN0 Ground

[DAQ_SYRS_369]

The AIN0 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #1.

#### PPAIO-16/4 ADC Module #1 Connection to AIN1

[DAQ_SYRS_370]

The first Adafruit ADS1115 module A1 input shall be connected to the positive side of the AIN1 screw terminal.

#### PPAIO-16/4 ADC Module #1 Connection to AIN1 Ground

[DAQ_SYRS_371]

The AIN1 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #1.

#### PPAIO-16/4 ADC Module #1 Connection to AIN2

[DAQ_SYRS_372]

The first Adafruit ADS1115 module A2 input shall be connected to the positive side of the AIN2 screw terminal.

#### PPAIO-16/4 ADC Module #1 Connection to AIN2 Ground

[DAQ_SYRS_373]

The AIN2 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #1.

#### PPAIO-16/4 ADC Module #1 Connection to AIN3

[DAQ_SYRS_374]

The first Adafruit ADS1115 module A3 input shall be connected to the positive side of the AIN3 screw terminal.

#### PPAIO-16/4 ADC Module #1 Connection to AIN3 Ground

[DAQ_SYRS_375]

The AIN3 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #1.

#### PPAIO-16/4 ADC Module #2 Address Selection

[DAQ_SYRS_376]

The first Adafruit ADS1115 module shall be assigned address one.

Note: Address zero is obtained by tying the ADDR pin to +5V.

#### PPAIO-16/4 ADC Module #2 Connection to AIN0

[DAQ_SYRS_377]

The first Adafruit ADS1115 module A0 input shall be connected to the positive side of the AIN4 screw terminal.

#### PPAIO-16/4 ADC Module #2 Connection to AIN4 Ground

[DAQ_SYRS_378]

The AIN4 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #2.

#### PPAIO-16/4 ADC Module #2 Connection to AIN5

[DAQ_SYRS_379]

The first Adafruit ADS1115 module A1 input shall be connected to the positive side of the AIN5 screw terminal.

#### PPAIO-16/4 ADC Module #2 Connection to AIN5 Ground

[DAQ_SYRS_380]

The AIN5 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #2.

#### PPAIO-16/4 ADC Module #2 Connection to AIN6

[DAQ_SYRS_381]

The first Adafruit ADS1115 module A2 input shall be connected to the positive side of the AIN6 screw terminal.

#### PPAIO-16/4 ADC Module #2 Connection to AIN6 Ground

[DAQ_SYRS_382]

The AIN6 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #2.

#### PPAIO-16/4 ADC Module #2 Connection to AIN7

[DAQ_SYRS_383]

The first Adafruit ADS1115 module A3 input shall be connected to the positive side of the AIN7 screw terminal.

#### PPAIO-16/4 ADC Module #2 Connection to AIN7 Ground

[DAQ_SYRS_384]

The AIN7 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #2.

#### PPAIO-16/4 ADC Module #3 Address Selection

[DAQ_SYRS_385]

The first Adafruit ADS1115 module shall be assigned address two.

Note: Address two is obtained by tying the ADDR pin to SDA on ADC Module #3.

#### PPAIO-16/4 ADC Module #3 Connection to AIN8

[DAQ_SYRS_386]

The first Adafruit ADS1115 module A0 input shall be connected to  the positive side of the AIN8 screw terminal.

#### PPAIO-16/4 ADC Module #3 Connection to AIN8 Ground

[DAQ_SYRS_387]

The AIN8 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #3.

#### PPAIO-16/4 ADC Module #3 Connection to AIN9

[DAQ_SYRS_388]

The first Adafruit ADS1115 module A1 input shall be connected to the positive side of the AIN9 screw terminal.

#### PPAIO-16/4 ADC Module #3 Connection to AIN9 Ground

[DAQ_SYRS_389]

The AIN9 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #3.

#### PPAIO-16/4 ADC Module #3 Connection to AIN10

[DAQ_SYRS_390]

The first Adafruit ADS1115 module A2 input shall be connected to the positive side of the AIN10 screw terminal.

#### PPAIO-16/4 ADC Module #3 Connection to AIN10 Ground

[DAQ_SYRS_391]

The AIN10 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #3.

#### PPAIO-16/4 ADC Module #3 Connection to AIN11

[DAQ_SYRS_392]

The first Adafruit ADS1115 module A3 input shall be connected to the positive side of the AIN11 screw terminal.

#### PPAIO-16/4 ADC Module #3 Connection to AIN11 Ground

[DAQ_SYRS_393]

The AIN11 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #3.

#### PPAIO-16/4 ADC Module #4 Address Selection

[DAQ_SYRS_394]

The first Adafruit ADS1115 module shall be assigned address three.

Note: Address three is obtained by tying the ADDR pin to SCL on ADC Module #4.

#### PPAIO-16/4 ADC Module #4 Connection to AIN12

[DAQ_SYRS_395]

The first Adafruit ADS1115 module A0 input shall be connected to  the positive side of the AIN12 screw terminal.

#### PPAIO-16/4 ADC Module #4 Connection to AIN2 Ground

[DAQ_SYRS_396]

The AIN12 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #4.

#### PPAIO-16/4 ADC Module #4 Connection to AIN13

[DAQ_SYRS_397]

The first Adafruit ADS1115 module A1 input shall be connected to the positive side of the AIN13 screw terminal.

#### PPAIO-16/4 ADC Module #4 Connection to AIN13 Ground

[DAQ_SYRS_398]

The AIN13 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #4.

#### PPAIO-16/4 ADC Module #4 Connection to AIN14

[DAQ_SYRS_399]

The first Adafruit ADS1115 module A2 input shall be connected to the positive side of the AIN14 screw terminal.

#### PPAIO-16/4 ADC Module #4 Connection to AIN14 Ground

[DAQ_SYRS_400]

The AIN14 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #4.

#### PPAIO-16/4 ADC Module #4 Connection to AIN15

[DAQ_SYRS_401]

The first Adafruit ADS1115 module A3 input shall be connected to the positive side of the AIN15 screw terminal.

#### PPAIO-16/4 ADC Module #4 Connection to AIN15 Ground

[DAQ_SYRS_402]

The AIN15 screw terminal Gnd pin shall be connected to the analog ground associated with ADC module #4.

#### PPAIO-16/4 DAC I2C Module Addresses (Daisy-Chaining)

[DAQ_SYRS_403]

When daisy-chaining two PPAIO-16/4, one board shall be populated with Adafruit MCP4725 DAC breakout boards and the second shall be populated with Sparkfun MCP4725 breakout boards.

The MCP4725 DAC chip (used on the Adafruit and Sparkfun MCP4725 breakout boards) only support two programmable addresses. Fortunately, the chips used by Adafruit and Sparkfun have different base I2C addresses (Adafruit = 0x62/0x63 and Sparkfun = 0x60/0x61). Therefore, by using breakout boards from the two different vendors it is possible to place 8 DACs on two I2C busses.

#### PPAIO-16/4 DAC Selection on First PPAIO-16/4

[DAQ_SYRS_404]

The first board in a pair of daisy-chained PPAIO-16/4 boards shall be populated with four Adafruit MCP4725 breakout boards.

#### PPAIO-16/4 DAC Selection on Second (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_405]

The second board in a pair of daisy-chained PPAIO-16/4 boards shall be populated with four Sparkfun MCP4725 breakout boards.

#### PPAIO-16/4 DAC #0 I2C Bus Selection on First (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_406]

DAC #0 (AOUT0) shall be connected to the "A" I2C bus on the PPAIO-16/4 Digital Input Connection (see [DAQ_SYRS_362]).

#### PPAIO-16/4 DAC #0 I2C Address Selection on First (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_407]

Adafruit DAC #0 (AOUT0) shall be have the ADDR pin tied to Gnd (I2C address=0x62, this is the default if the pin is left floating).

#### PPAIO-16/4 DAC #1 I2C Bus Selection on First (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_408]

DAC #1 (AOUT1) shall be connected to the "A" I2C bus on the PPAIO-16/4 Digital Input Connection (see [DAQ_SYRS_362]).

#### PPAIO-16/4 DAC #1 I2C Bus Selection on First (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_409]

Adafruit DAC #1 (AOUT1) shall be have the ADDR pin tied to Vcc (I2C address=0x63).

#### PPAIO-16/4 DAC #2 I2C Bus Selection on First (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_410]

DAC #2 (AOUT2) shall be connected to the "B" I2C bus on the PPAIO-16/4 Digital Input Connection (see [DAQ_SYRS_362]).

#### PPAIO-16/4 DAC #2 I2C Address Selection on First (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_411]

Adafruit DAC #2 (AOUT2) shall be have the ADDR pin tied to Gnd (I2C address=0x62, this is the default if the pin is left floating).

#### PPAIO-16/4 DAC #3 I2C Bus Selection on First (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_412]

DAC #3 (AOUT3) shall be connected to the "B" I2C bus on the PPAIO-16/4 Digital Input Connection (see [DAQ_SYRS_362]).

#### PPAIO-16/4 DAC #3 I2C Bus Selection on First (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_413]

Adafruit DAC #3 (AOUT3) shall be have the ADDR pin tied to Vcc (I2C address=0x63).

#### PPAIO-16/4 DAC #0 I2C Bus Selection on Second (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_414]

DAC #0 (AOUT0) shall be connected to the "A" I2C bus on the PPAIO-16/4 Digital Input Connection (see [DAQ_SYRS_362]).

#### PPAIO-16/4 DAC #0 I2C Address Selection on Second (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_415]

Sparkfun DAC #0 (AOUT0) shall be have the A0 pin tied to Gnd (I2C address=0x60, this is the default if the pin is left floating).

#### PPAIO-16/4 DAC #1 I2C Bus Selection on Second (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_416]

DAC #1 (AOUT1) shall be connected to the "A" I2C bus on the PPAIO-16/4 Digital Input Connection (see [DAQ_SYRS_362]).

#### PPAIO-16/4 DAC #1 I2C Bus Selection on Second (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_417]

Sparkfun DAC #1 (AOUT1) shall be have the A0 pin tied to Vcc (I2C address=0x61).

#### PPAIO-16/4 DAC #2 I2C Bus Selection on Second (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_418]

DAC #2 (AOUT2) shall be connected to the "B" I2C bus on the PPAIO-16/4 Digital Input Connection (see [DAQ_SYRS_362]).

#### PPAIO-16/4 DAC #2 I2C Address Selection on Second (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_419]

Sparkfun DAC #2 (AOUT2) shall be have the A0 pin tied to Gnd (I2C address=0x60, this is the default if the pin is left floating).

#### PPAIO-16/4 DAC #3 I2C Bus Selection on Second (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_420]

DAC #3 (AOUT3) shall be connected to the "B" I2C bus on the PPAIO-16/4 Digital Input Connection (see [DAQ_SYRS_362]).

#### PPAIO-16/4 DAC #3 I2C Bus Selection on Second (in a daisy chain) PPAIO-16/4

[DAQ_SYRS_421]

Sparkfun DAC #3 (AOUT3) shall be have the A0 pin tied to Vcc (I2C address=0x61).

#### PPAIO-16/4 Power Supply Connection

[DAQ_SYRS_422]

The PPAIO-16/4 board shall provide a two-pin screw terminal connector (5mm centers) to accept ±12V/Gnd from an external power supply

#### PPAIO-16/4 Clean Power Supply

[DAQ_SYRS_423]

The power supply connected to the PPAIO-16/4 shall be a "clean" analog-only power supply that does not power digital boards or other noisy devices.

#### PPAIO-16/4 Power Supply LED

[DAQ_SYRS_424]

The PPAIO-16/4board shall provide an LED that is illuminated whenever power is applied to the board.

#### PPAIO-16/4 Output Voltage Conversion

[DAQ_SYRS_425]

The PPAIO-16/4board shall provide four analog amplifier channels to convert the 0-5V output from the DAC to the range ±5V (or ±10V).

### 4-Channel Analog Conditioning Board (PPAC-4)

#### PPAC-4 Capacity

[DAQ_SYRS_450]

The PPAC-4 board shall provide 4 channels of analog conditioning.

#### PPAC-4 Power Supply Connection

[DAQ_SYRS_451]

The PPAC-4 board shall provide a three-pin screw terminal connector (5mm centers) to accept ±12V/Gnd from an external power supply

#### PPAC-4 Power Supply LED

[DAQ_SYRS_452]

The PPAC-4 board shall provide an LED that is illuminated whenever power is applied to the board (+12V).

#### PPAC-4 Clean Power Supply

[DAQ_SYRS_453]

The power supply connected to the PPAC-4 shall be a "clean" analog-only power supply that does not power digital boards or other noisy devices.

#### PPAC-4 Signal Isolation

[DAQ_SYRS_454]

Each analog input channel shall provide an isolation amplifier that isolates the analog input signal from the DAQ boards as well as isolating the input signal from other analog signals.

#### PPAC-4 Power Supply Isolation

[DAQ_SYRS_455]

Each analog input channel shall provide an isolated power supply for each isolation amplifier.

#### PPAC-4 Maximum Input Voltage Range

[DAQ_SYRS_456]

Each analog input channel shall accept inputs in the range ±10V.

#### PPAC-4 Double-ended Output

[DAQ_SYRS_457]

Each analog channel shall produce a pair of double-ended outputs, whose output voltage is in the range 0-4.096V.

#### PPAC-4 Double-ended Output from Positive/Zero Inputs

[DAQ_SYRS_458]

Each analog channel shall produce a positive (or zero) voltage on the positive terminal that is greater than the positive (or zero) output on the negative terminal.

Note: for double-ended outputs, the resulting value is the difference between the positive and negative output terminals – both of which are positive voltages.

#### PPAC-4 Double-ended Output from Negative Inputs

[DAQ_SYRS_459]

Each analog channel shall produce a positive (or zero) voltage on the positive terminal that is less than the positive (or zero) output on the negative terminal.

Note: for double-ended outputs, the resulting value is the difference between the positive and negative output terminals – both of which are positive voltages.

#### PPAC-4 Linear Conversion

[DAQ_SYRS_460]

The transfer function for the analog conversion shall be linear.

#### PPAC-4 Span/Gain Adjustment

[DAQ_SYRS_461]

The PPAC-4 board shall provide a multi-turn potentiometer (on each channel) that allows a technician to adjust the gain (span) of an input signal.

Note: the span adjustment is intended for minor gain tweaks when calibrating the device at site. The Adafruit ADS1115 module provides a programmable gain amplifier than the software can use to handle larger gain adjustments (such as ±5V inputs rather than ±10V) to avoid losing resolution in the ADC.

#### PPAC-4 Zero Adjustment

[DAQ_SYRS_462]

The PPAC-4 board shall provide a multi-turn potentiometer (on each channel) that allows a technician to adjust the zero point (offset) of an input signal.

#### PPAC-4 VRef Adjustment

[DAQ_SYRS_463]

The PPAC-4 board shall provide a multi-turn potentiometer (on each channel) that allows a technician to adjust the voltage reference point (the cross-over point between the two legs of the differential/doubled-ended output).

Note: This is usually set to ½ of the maximum output voltage (4.096V is the typical maximum output voltage, so VRef is typically set to 2.048V).

#### PPAC-4 Analog Input Terminals

[DAQ_SYRS_464]

The PPAC-4 board shall provide 4 two-pin screw terminals to connect an input voltage to the input pins on each channel's isolation amplifier.

#### PPAC-4 Single-ended Inputs

[DAQ_SYRS_465]

Each input channel on the PPAC-4 board shall be a single-ended input (voltage input + Gnd).

#### PPAC-4 Analog Output Terminals

[DAQ_SYRS_466]

The PPAC-4 board shall provide 4 two-pin screw terminals to connect an input voltage to the input pins on each channel's isolation amplifier.

#### PPAC-4 Frequency Response

[DAQ_SYRS_467]

Each channel on the PPAC-4 board shall have a frequency response of at least 1kHz.

### 12-Channel Relay-based Digital I/O Board (PPRlyio-12)

#### PPRLYIO-12 Capacity

[DAQ_SYRS_500]

The PPRLYIO-12 shall provide 12 bits of relay digital I/O.

#### PPRLYIO-12 Power Supply Connection

[DAQ_SYRS_501]

The PPRLYIO-12  board shall provide a two-pin screw terminal connector (5mm centers) to accept 5V/Gnd from an external power supply

#### PPRLYIO-12 Power Supply LED

[DAQ_SYRS_502]

The PPRLYIO-12  board shall provide an LED that is illuminated whenever power is applied to the board.

#### PPRLYIO-12 Output Connector Pinout

[DAQ_SYRS_503]

The PPRLYIO-12 board shall provide a 20-pin (2x10) header that is compatible with the PPDIO96 bank pinouts, using the following pinout:

- Bit 1

- Bit 0

- Bit 2

- Gnd

- Bit 3

- Gnd

- Bit 4

- Gnd

- Bit 5

- Gnd

- Bit 6

- Gnd

- Bit 7

- Gnd

- Bit 8

- Gnd

- Bit 9

- Gnd

- Bit 10

- Bit 11

Note: Gnd pins are all connected to the same ground plane.

#### PPRLYIO-12 Input Connector Pinout

[DAQ_SYRS_504]

The PPRLYIO-12 board shall provide a 20-pin (2x10) header that is compatible with the PPDIO96 bank pinouts, using the same pinout as in [DAQ_SYRS_503].

#### PPRLYIO-12 Signal LEDs

[DAQ_SYRS_505]

The PPRLYIO-12 board shall provide 12 signal LEDs.

#### PPRLYIO-12 Signal LED #0 Illumination

[DAQ_SYRS_506]

Signal LED #0 on the PPRLYIO-12 board shall illuminate when input D0 is active.

#### PPRLYIO-12 Signal LED #1 Illumination

[DAQ_SYRS_507

Signal LED #1 on the PPRLYIO-12 board shall illuminate when input D1 is active.

#### PPRLYIO-12 Signal LED #2 Illumination

[DAQ_SYRS_508]

Signal LED #2 on the PPRLYIO-12 board shall illuminate when input D2 is active.

#### PPRLYIO-12 Signal LED #3 Illumination

[DAQ_SYRS_509]

Signal LED #3 on the PPRLYIO-12 board shall illuminate when input D3 is active.

#### PPRLYIO-12 Signal LED #4 Illumination

[DAQ_SYRS_510]

Signal LED #4 on the PPRLYIO-12 board shall illuminate when input D4 is active.

#### PPRLYIO-12 Signal LED #5 Illumination

[DAQ_SYRS_511]

Signal LED #5 on the PPRLYIO-12 board shall illuminate when input D5 is active.

#### PPRLYIO-12 Signal LED #6 Illumination

[DAQ_SYRS_512]

Signal LED #6 on the PPRLYIO-12 board shall illuminate when input D6 is active.

#### PPRLYIO-12 Signal LED #7 Illumination

[DAQ_SYRS_213]

Signal LED #7 on the PPRLYIO-12 board shall illuminate when input D7 is active.

#### PPRLYIO-12 Signal LED #8 Illumination

[DAQ_SYRS_514]

Signal LED #8 on the PPRLYIO-12 board shall illuminate when input D8 is active.

#### PPRLYIO-12 Signal LED #9 Illumination

[DAQ_SYRS_515]

Signal LED #9 on the PPRLYIO-12 board shall illuminate when input D9 is active.

#### PPRLYIO-12 Signal LED #10 Illumination

[DAQ_SYRS_516]

Signal LED #10 on the PPRLYIO-12 board shall illuminate when input D10 is active.

#### PPRLYIO-12 Signal LED #11 Illumination

[DAQ_SYRS_517]

Signal LED #11 on the PPRLYIO-12 board shall illuminate when input D11 is active.

#### PPRLYIO-12 Input Polarity Selection

[DAQ_SYRS_518]

Each input channel on the PPRlyio-12 board shall provide a jumper to select active high or active low operation.

#### PPRLYIO-12 Output Pullup Resistor Selection

[DAQ_SYRS_519]

Each output channel on the PPRlyio-12 board shall provide a jumper to open collector or TTL-level (via pullup resistor) operation.

#### PPRELAY12 Output Isolation

[DAQ_SYRS_520]

The PPRlyio-12 board shall provide isolation for all relay outputs.

Note: this feature is provided by default given the nature of mechanical relays.

#### PPRELAY12 Output Current Capacity

[DAQ_SYRS_521]

The PPRlyio-12 board shall be capable of handling an minimum of 2A on each output channel.

#### PPRELAY12 Output Voltage Capacity

[DAQ_SYRS_522]

The PPRlyio-12 board shall be capable of handling a minimum of  30VDC on each output channel.

#### PPRELAY12 Output Connectors

[DAQ_SYRS_523]

The PPRlyio-12 board shall provide a three-pin screw terminal for each relay output (5mm centers) for NC/COMM/NO relay connections.

#### PPRELAY12 NC Optional

[DAQ_SYRS_524]

The PPRlyio-12 board shall provide the option of using SPST relays (with NO/COM connections only) by swapping relays and substituting two-terminal connections in place of the three-terminal connections.

#### PPRELAY12 Output Control

[DAQ_SYRS_525]

The PPRlyio-12 board shall buffer (and possibly invert, depending on the state of the corresponding polarity jumper setting) the input signal on each DIN0-11 pin and pass it through to the corresponding DOUT0-11 pin.

### PPAC420 4-20mA Analog Input Conditioning Board (PPAC420)

#### PPAC420  Capacity

[DAQ_SYRS_550]

The PPAC420 board shall provide 8 channels of analog conditioning.

#### PPAC420  Power Supply Connection

[DAQ_SYRS_551]

The PPAC420  board shall provide a two-pin screw terminal connector (5mm centers) to accept 5V/Gnd from an external power supply.

#### PPAC420 Power Supply LED

[DAQ_SYRS_552]

The PPAC420 board shall provide an LED that is illuminated whenever power is applied to the board (+5V).

#### PPAC420 Clean Power Supply

[DAQ_SYRS_553]

The power supply connected to the PPAC420 shall be a "clean" analog-only power supply that does not power digital boards or other noisy devices.

#### PPAC420 Maximum Input Current Range

[DAQ_SYRS_554]

Each analog input channel shall accept currents in the range 0-20mA.

#### PPAC420 Output Voltage Range

[DAQ_SYRS_555]

Each analog channel shall produce single-ended outputs in the range -1.25V to +5V (-1.25V = 0ma, 0V=4mA, 5V=20mA).

#### PPAC420 Analog Input Terminals

[DAQ_SYRS_556]

The PPAC420 board shall provide 8 two-pin screw terminals to connect a current loop to the inputs on the board.

#### PPAC420 Analog Output Terminals

[DAQ_SYRS_557]

The PPAC420 board shall provide 8 two-pin screw terminals to connect the output voltage to some external device (such as the PPAC4).

## Software Requirements

### DIS64 Emulation

DIS64 emulation requirements are deprecated at this point.

#### DIS64 CPU Interface Board Usage

[DAQ_SYRS_600]

The DIS64 Software shall use one DAQIF board with a Teensy 3.2 CPU module installed.

#### DIS64 PPDIO96 Parallel I/O Board Usage

[DAQ_SYRS_601]

The DIS64 Software shall use one PPDIO96 board.

#### DIS64 Communications Channel

[DAQ_SYRS_602]

The DIS64 Software shall communicate with the host system using RS-232 communications.

#### DIS64 Communication Speed

[DAQ_SYRS_603]

The DIS64 Software shall communicate with the host system at 9600 bps.

#### DIS64 Communication Stop Bits

[DAQ_SYRS_604]

The DIS64 Software shall require only one stop bit during serial communication.

#### DIS64 Communication Parity

[DAQ_SYRS_605]

The DIS64 Software shall not provide (nor require) any parity information during serial communication.

#### DIS64 Input Debouncing

[DAQ_SYRS_606]

The DIS64 Software use a 20-30ms debounce period for all digital inputs.

#### DIS64 Number of Supported Digital Inputs

[DAQ_SYRS_607]

The DIS64 Software shall support 64 digital inputs (D0-D63 on the PPDIO96 board).

Note: Although the DIS64 software could technically read all 96 input bits and make them available to the host machine (the underlying protocol could also support 96 inputs), the DIS64 software exists to support legacy applications, none of which would be able to handle more than 64 inputs.

#### DIS64 Unsupported Digital Inputs

[DAQ_SYRS_608]

The DIS64 Software shall ignore inputs D64-D95 on the PPDIO96 board.

#### DIS64 I'm Alive

[DAQ_SYRS_609]

The DIS64 Software shall transmit a single byte (containing 0 or 128) to the host every 5 seconds (±0.5 secs).

#### DIS64 I'm Alive and Poked

[DAQ_SYRS_610]

The DIS64 Software shall transmit a single byte containing 0 to the host every 5 seconds (±0.5 secs) after receiving "All On Channels" (0xfe) from the host.

#### DIS64 I'm Alive and Virgin

[DAQ_SYRS_611]

After reset or power one, the DIS64 Software shall transmit a single byte containing 128 to the host every 5 seconds (±0.5 secs) until receiving "All On Channels" (0xfe) from the host.

#### DIS64 Data Transmission

[DAQ_SYRS_612]

The DIS64 Software shall transmit a single byte when a digital input pin changes state.

#### DIS64 Data Transmission on Digital Input High Transition

[DAQ_SYRS_613]

The DIS64 Software shall transmit a single byte containing a single byte with the H.O. bit set whenever there is a 0-to-1 transition (open-to-close) on one of the 64 digital inputs.

#### DIS64 Data Transmission on Digital Input Low Transition

[DAQ_SYRS_614]

The DIS64 Software shall transmit a single byte containing a single byte with the H.O. bit clear whenever there is a 1-to-0 transition (close-to-open) on one of the 64 digital inputs.

#### DIS64 Data Transmission Value on Digital Input Transition

[DAQ_SYRS_615]

The DIS64 Software shall transmit the digital input port number, plus one, in the L.O. seven bits whenever a transition takes place on one of the digital inputs (e.g, a transition on port 10 will transmit the value 0xB in the L.O. seven bits of the transmitted byte).

#### DIS64 Request for All Channels On command

[DAQ_SYRS_616]

The DIS64 Software shall recognize a single byte value, 0xFE, from the host computer. This value instructs the DIS64 software to switch to "I'm Alive and Poked" mode (versus "I'm Alive and Virgin" which is the default mode on power up). See [DAQ_SYRS_611] and [DAQ_SYRS_611] for more details.

#### DIS64 Ignore All Serial Input Characters Except 0xFE

[DAQ_SYRS_617]

The DIS64 Software shall ignore all characters received on the serial port except the "Request for All Channels On" command (see [DAQ_SYRS_616]).

### PPDAQ Standard Control Software

#### PPDAQ Standard Software Platform

[DAQ_SYRS_700]

The PPDAQ software shall run on a Netburner MOD45515 MOD-70 evaluation board connected to a DAQIF interface board.

#### PPDAQ Communications Protocol – RS-232

[DAQ_SYRS_701]

The PPDAQ software shall accept commands from a host computer via the UART1 RS-232 serial port on the Netburner MOD45515 MOD-70 evaluation board.

#### PPDAQ RS-232 Enabled

[DAQ_SYRS_702]

The PPDAQ software shall activate the RS-232 communications protocol if DIP switch #1 on the Netburner MOD45515 MOD-70 evaluation board is set to the on position.

#### PPDAQ RS-232 Disabled

[DAQ_SYRS_703]

The PPDAQ software shall not activate the RS-232 communications protocol if DIP switch #1 on the Netburner MOD45515 MOD-70 evaluation board is set to the off position.

#### PPDAQ Communications Protocol – USB

[DAQ_SYRS_704]

The PPDAQ software shall accept commands from a host computer via the USB port on the Netburner MOD45515 MOD-70 evaluation board.

#### PPDAQ USB Enabled

[DAQ_SYRS_705]

The PPDAQ software shall activate the USB communications protocol if DIP switch #2 on the Netburner MOD45515 MOD-70 evaluation board is set to the on position.

#### PPDAQ USB Disabled

[DAQ_SYRS_706]

The PPDAQ software shall not activate the USB communications protocol if DIP switch #2 on the Netburner MOD45515 MOD-70 evaluation board is set to the off position.

Note: to the PPDAQ software, the Netburner MOD45515 MOD-70 evaluation board USB port appears as a serial/COM port.

#### PPDAQ Communications Protocol – Ethernet

[DAQ_SYRS_707]

The PPDAQ software shall accept commands from a host computer via the Ethernet port on the Netburner MOD45515 MOD-70 evaluation board.

#### PPDAQ Ethernet IP Address

[DAQ_SYRS_708]

The PPDAQ software shall set the Ethernet IP address to a value in the range 192.168.2.70-192.168.2.73 based on DIP switch 5-6 settings on the Netburner.

#### PPDAQ Ethernet IP Address 192.168.2.70

[DAQ_SYRS_709

]

The PPDAQ software shall set the Ethernet IP address to 192.168.2.70 if the Netburner dip switches 5-6 are set to (off, off).

#### PPDAQ Ethernet IP Address 192.168.2.71

[DAQ_SYRS_710]

The PPDAQ software shall set the Ethernet IP address to 192.168.2.71 if the Netburner dip switches 5-6 are set to (on, off).

#### PPDAQ Ethernet IP Address 192.168.2.72

[DAQ_SYRS_711]

The PPDAQ software shall set the Ethernet IP address to 192.168.2.72 if the Netburner dip switches 5-6 are set to (off, on).

#### PPDAQ Ethernet IP Address 192.168.2.73

[DAQ_SYRS_712]

The PPDAQ software shall set the Ethernet IP address to 192.168.2.73 if the Netburner dip switches 5-6 are set to (on, on).

#### PPDAQ Ethernet IP Address 192.168.2.74

[DAQ_SYRS_713]

(Deprecated)

The PPDAQ software shall set the Ethernet IP address to 192.168.2.74 if the Netburner DIP switches 1-3 are set to (off, off, on).

#### PPDAQ Ethernet IP Address 192.168.2.75

[DAQ_SYRS_714]

(Deprecated)

The PPDAQ software shall set the Ethernet IP address to 192.168.2.75 if the Netburner DIP switches 1-3 are set to (on, off, on).

#### PPDAQ Ethernet IP Address 192.168.2.76

[DAQ_SYRS_715]

(Deprecated)

The PPDAQ software shall set the Ethernet IP address to 192.168.2.76 if the Netburner DIP switches 1-3 are set to (off, on, on).

#### PPDAQ Ethernet Enabled

[DAQ_SYRS_716]

The PPDAQ software shall enable Ethernet operation if the Netburner DIP switch 3 is in the on position.

#### PPDAQ Ethernet Disabled

[DAQ_SYRS_716.5]

The PPDAQ software shall disable Ethernet operation if the Netburner DIP switch 3 is in the off position.

#### PPDAQ Ethernet Port

[DAQ_SYRS_717]

The PPDAQ software shall communicate via Ethernet using socket port 0x5050 (decimal 20560, "PP", for "Plantation Productions").

#### PPDAQ Ethernet Multiple Clients enabled

[DAQ_SYRS_718]

The PPDAQ software shall allow up to 5 Ethernet clients if the Netburner DIP switch #4 is set to the on position.

#### PPDAQ Unit Test Mode Enabled

[DAQ_SYRS_719]

The PPDAQ software shall operate in a special "unit test" mode if the Netburner dip switch #7 is set to the on position.

#### PPDAQ Unit Test Mode Disabled

[DAQ_SYRS_720]

The PPDAQ software shall operate in the normal (non-unit-test) mode if the Netburner dip switch #7 is set to the off position.

#### PPDAQ Unit Test Mode I/O

[DAQ_SYRS_721]

The PPDAQ software shall utilize the UART0 serial port on the Netburner MOD45515 MOD-70 evaluation board for unit test communication.

#### PPDAQ Application-Defined DIP Switch

[DAQ_SYRS_722]

(Deprecated)

The PPDAQ software shall reserve the use of the Netburner DIP switch #7 for host computer/user application purposes.

Note: the PPDIO96 boards also have 8 dip switches that are reserved for host application use.

#### PPDAQ Read DIP Switches

[DAQ_SYRS_723]

The PPDAQ software shall allow the host software to read the current state of the DIP switches on the Netburner MOD45515 MOD-70 evaluation board.

#### PPDAQ Event Log

[DAQ_SYRS_724]

The PPDAQ software shall write event log entries to the Netburner MOD45515 MOD-70 evaluation board UART0 RS-232 serial port.

#### PPDAQ Command Protocol

[DAQ_SYRS_725]

The PPDAQ software shall wait for a command from the host, process that command, and respond to the host.

Note: the PPDAQ software is strictly command-driven. It does not initiate and communication on its own, it only responds to commands sent from the host system.

#### PPDAQ Command Source #1

[DAQ_SYRS_726]

The PPDAQ software shall accept commands from the UART1 serial port on the Netburner MOD45515 MOD-70 evaluation board if serial communications are enabled.

#### PPDAQ Command Source #2

[DAQ_SYRS_727]

The PPDAQ software shall accept commands from the USB port on the Netburner MOD45515 MOD-70 evaluation board if USB communications are enabled.

#### PPDAQ Command Source #3

[DAQ_SYRS_728]

The PPDAQ software shall accept commands from the Ethernet port on the Netburner MOD45515 MOD-70 evaluation board if Ethernet communications are enabled.

#### PPDAQ Command Format

[DAQ_SYRS_729]

The PPDAQ software shall processing commands consisting of lines of ASCII text terminated by a new-line character ('\n', 0xA, line-feed). An optional carriage return ("\r", 0xD) may precede the line feed.

#### PPDAQ Command Syntax

[DAQ_SYRS_730]

The PPDAQ command syntax shall take the following basic form:

<command>  <subcommand>  <arguments>

Note: the syntax for the arguments depends on the particular command. Some commands may not allow any arguments or even subcommands, some may require a fixed number of arguments, some may allow a variable number of arguments. Whitespace (spaces or tabs) must separate commands from arguments and arguments from one another.

#### PPDAQ Command Case Insensitive

[DAQ_SYRS_731]

PPDAQ commands shall be case insensitive.

Note: all examples in this document will use lower case alphabetic characters. This does not imply that lower case is required.

#### PPDAQ Whitespace Ignored

[DAQ_SYRS_732]

PPDAQ commands shall allow zero or more white space characters (spaces and tabs) around any command or argument.

#### PPDAQ Connect Command Syntax

[DAQ_SYRS_733]

(Deprecated)

The PPDAQ connect command shall take the following form:

connect\n

After a host is physically connected, it must issue the connect command to inform the PPDAQ that it is going to communicate with the PPDAQ system.

#### PPDAQ Connect Response

[DAQ_SYRS_734]

(Deprecated)

The PPDAQ software shall return the original command line as the response if the connection is valid.

#### PPDAQ At Most One Serial Connection

[DAQ_SYRS_735]

(Deprecated)

The PPDAQ software shall only recognize a single connection on the RS-232 port.

Note: if the host software reissues the connect command on the RS-232 port without first disconnecting, the PPDAQ software shall respond with the following error response:

Error: RS232 already connected: <original connection command line> \n

#### PPDAQ At Most One USB Connection

[DAQ_SYRS_736]

(Deprecated)

The PPDAQ software shall only recognize a single connection on the USB port.

Note: if the host software reissues the connect command on the USB port without first disconnecting, the PPDAQ software shall respond with the following error response:

Error: USB already connected: <original connection command line> \n

#### PPDAQ Maximum Ethernet Connections #1

[DAQ_SYRS_737]

The PPDAQ software shall only recognize a single connection on the Ethernet port if the Netburner DIP switch #4 is in the off position.

Note: if the host software reissues the connect command on the Ethernet port without first disconnecting, the PPDAQ software shall respond with the following error response:

Error: Eth-Connection: <original connection command line> \n

#### PPDAQ Maximum Ethernet Connections #2

[DAQ_SYRS_738]

The PPDAQ software shall only recognize up to five connections on the Ethernet port if the Netburner DIP switch #4 is in the on position.

Note: if the host software reissues the connect command on the Ethernet port without first disconnecting at least one connection, the PPDAQ software shall respond with the following error response:

Error: Eth-Connection: <original connection command line> \n

#### PPDAQ Disconnect Command Syntax

[DAQ_SYRS_739]

(Deprecated)

The PPDAQ disconnect command shall take the following form:

disconnect\n

This terminates the communication stream between the host and the PPDAQ software on the communications port where the disconnect command was issued.

#### PPDAQ Disconnect Response

[DAQ_SYRS_740]

(Deprecated)

The PPDAQ software shall return the original command line as the response if the disconnect command is valid.

#### PPDAQ Disconnect Command While Not Connected

[DAQ_SYRS_741]

(Deprecated)

If the host issues a disconnect command on a port to which is has not connected via the connect command, the PPDAQ software shall return the following error response:

Error: Not connected: <original command line> \n

#### PPDAQ Watchdog Refresh While Running

[DAQ_SYRS_742]

The PPDAQ software shall refresh the DAQIF watchdog timer with a period of less than two seconds.

#### PPDAQ Watchdog Timeout While Disconnected

[DAQ_SYRS_743]

(Deprecated)

When no connection on any of the ports is made, the PPDAQ software shall not refresh the DAQIF watchdog timer.

Note: after 2-10 seconds without a connection, the watchdog timer will timeout. This will place the output relays on the PPSSR-16 and PPRELAY-12 boards in a fail-safe state (all relays off/open).

#### PPDAQ Reset Command Syntax

[DAQ_SYRS_744]

The PPDAQ reset command shall take the following form:

reset\n

This command will send a 350msec pulse on the DAQ reset line. In the existing design, a reset pulse will:

- Reset the watchdog timeout latch

- Reset all the MCP23S17 I/O expansion chips on the PPDIO96 boards (see the MCP23S17 data sheet for the implications of a reset).

- Clear all the PPSSR-16 and PPRELAY-12 output bits (set the relays to the open/fail-safe condition).

#### PPDAQ Version Command Syntax

[DAQ_SYRS_745]

The PPDAQ version command shall take the following form:

version\n

#### PPDAQ Reset Response

[DAQ_SYRS_746]

The PPDAQ software shall respond to the reset command by returning the command line sent to it.

#### PPDAQ Version Response

[DAQ_SYRS_747]

The PPDAQ version command shall return a string of the form:

PPDAQ:xx.yy\n

Where "xx" is a major version number and "yy" is a minor version number.

#### PPDAQ Echo Command Syntax

[DAQ_SYRS_748]

The PPDAQ echo command shall take the following form:

echo\n

#### PPDAQ Echo Response

[DAQ_SYRS_749]

The PPDAQ software shall respond to the echo command by returning the command line sent to it.

Note: main use of the echo command is to provide a "cleanup" command. If the command processor gets out of sync, sending an echo command and then waiting until you retrieve the corresponding echo can clean up the communication channel.

#### PPDAQ Help Command Syntax

[DAQ_SYRS_750]

The PPDAQ help command shall take the following form:

help\n

#### PPDAQ Help Response

[DAQ_SYRS_751]

The PPDAQ software shall respond to the help command by returning a string containing several lines of text containing a help message (contents of the help message is undefined by the requirements, to be determined later).

Note: the help command is intended for interactive usage via a terminal program, not by host software. Largely, it is intended for system testing and debugging purposes.

PPDAQ PPDIO96 and PPDI48 Support

##### PPDAQ PPDIO96/PPDI48 Board Count

[DAQ_SYRS_752]

The PPDAQ software shall support 0 to 6 PPDIO96 and PPDI48 boards.

##### PPDAQ PPDIO96 Bank Data Direction

[DAQ_SYRS_753]

The PPDAQ software shall allow individual banks (12 I/O lines) to be programmed as input or output.

##### PPDAQ PPDIO96 Default Bank Data Direction

[DAQ_SYRS_754]

The PPDAQ software shall default all PPDIO96 banks to input devices.

##### PPDAQ PPDIO96 Support

[DAQ_SYRS_755]

(Deprecated)

The PPDAQ software shall support 0 to 6 PPDIO96 boards.

##### PPDAQ PPDIO96 Interrupt Support

[DAQ_SYRS_756]

(Deprecated)

The PPDAQ software shall not support interrupt operation on the PPDIO96 board.

##### PPDAQ PPDIO96 Polarity Support

[DAQ_SYRS_757]

The PPDAQ software shall support programming the input polarity of individual input bits on the PPDIO96.

Note: Although the MCP23S17 I/O expander chip supports setting the polarity in hardware, there is no requirement that the PPDAQ software use this feature. I.e., the software could handle the polarity calculation without using the hardware.

##### PPDAQ PPDIO96 Pullup Support

[DAQ_SYRS_758]

The PPDAQ software shall support activating or deactivating the programmable pull-up resistors on the MCP23s17.

##### PPDAQ PPDIO96 Pullup Support #2

[DAQ_SYRS_759]

The PPDAQ software shall support activating pull-up resistors for individual input bits on the PPDIO96.

##### PPDAQ PPDIO96 Polling Rate

[DAQ_SYRS_760]

The PPDAQ software shall read all the digital inputs (on all PPDIO96 boards) every 25 – 100 msec (10Hz to 40 Hz).

##### PPDAQ PPDIO96 Input FIFO

[DAQ_SYRS_761]

The PPDAQ software shall maintain a 40-entry FIFO history list for each bit read from the PPDIO96.

##### PPDAQ PPDIO96 Input FIFO Clear

[DAQ_SYRS_762]

The PPDAQ software shall clear (empty) the PPDIO96 FIFO history list after the host computer reads the data.

##### PPDAQ PPDIO96 Input FIFO Overflow

[DAQ_SYRS_763]

The PPDAQ software shall only keep the last 40 readings in the FIFO queue while waiting for the host to read the data.  If more than 40 readings are taken while waiting for a host request, the PPDAQ software shall delete the oldest (greater than 40 readings) entries.

##### PPDAQ PPDIO96 Input FIFO Filtering

[DAQ_SYRS_764]

The PPDAQ software shall allow five different filtering options on PPDIO96 input data:

[DAQ_SYRS_765]

- Most recent reading in history list (last value read from the PPDIO96)

[DAQ_SYRS_766]

- First reading in history list (typically the first element read after the last host data request)

[DAQ_SYRS_767]

- Vote: the returned value is the largest number of zeros or ones (if there is a tie, the most recent reading wins).

[DAQ_SYRS_768

- Loser: the returned value is the lesser number of zeros or ones (if there is a tie, the most recent reading wins).

[DAQ_SYRS_769]

- Debounce: the returned value is the first sequence, looking backwards in the list, with a minimum length specified by the host computer. If no such sequence exists, the System returns the previous debounced value.

Note: debounce times are multiples of the polling rate (i.e., 25-100 msec).

##### PPDAQ PPDIO96 Setting Filtering

[DAQ_SYRS_770]

The PPDAQ software shall allow the host system to set the input filtering for individual bits.

Note: filtering is specified by (board, bank, bit, filter_type) where filter type is

- 0: Most recent reading

- 1: First reading in history list

- 2: Vote

- 3: Loser

- 4: Debounce

##### PPDAQ PPDIO96 Debounce Time Selection

[DAQ_SYRS_771]

The PPDAQ software shall allow the host to set the debounce time for a given input bit that uses debounce filtering.

Note: debounce time is set by (board, bank, bit, time) where time is the number of 25-100 msec time periods for the debounce operation.

Note: if the specified bit is not "debounce" filtering type, the sysem will save the debounce time and use it as the debounce value when the bit is switched to debounce filtering.

##### PPDAQ PPDIO96 Default Debounce Time

[DAQ_SYRS_772]

The PPDAQ software shall use one polling period (25-100 msec) as the default debounce time.

##### PPDAQ PPDIO96 Maximum Debounce Time

[DAQ_SYRS_773]

The PPDAQ software shall limit the debounce time to 40 polling cycles (1 sec).

Note: if the host software specifies a debounce time outside the range 1-40, the PPDAQ software shall ignore the value.

##### PPDAQ PPDIO96 Read Debounce Time

[DAQ_SYRS_774]

The PPDAQ software shall allow the host computer to read the debounce time for a particular bit.

Note: desired bit is specified by (board, bank, bit). PPDAQ returns an unsigned integer with the specified debounce time.

##### PPDAQ PPDIO96 Read Bit

[DAQ_SYRS_775]

The PPDAQ software shall allow the host to read a particular bit from the PPDIO96.

Note: bit to read is specified by (board, bank, bit).

Note: this operation reads data from internal PPDAQ memory, it doesn't force the system to read the actual hardware bits.

Note: bits read reflect the filter type associated with each bit.

Note: if specified bit is an output bit, the value read is undefined.

##### PPDAQ PPDIO96 Read Bank

[DAQ_SYRS_776]

The PPDAQ software shall allow the host to read a particular bank (12 bits) from the PPDIO96.

Note: bank to read is specified by (board, bank).

Note: this operation reads data from internal PPDAQ memory, it doesn't force the system to read the actual hardware bits.

Note: bits read reflect the filter type associated with each bit.

Note: if specified bank is an output bank, the value read is undefined.

##### PPDAQ PPDIO96 Read Board

(Deprecated)

[DAQ_SYRS_777]

The PPDAQ software shall allow the host to read a particular board (96 bits) from the PPDIO96.

Note: board to read is specified by (board).

Note: this operation reads data from internal PPDAQ memory, it doesn't force the system to read the actual hardware bits.

Note: bits read reflect the filter type associated with each bit.

Note: if specified board contains an output bank, the value read is undefined.

##### PPDAQ PPDIO96 Read All

(Deprecated)

[DAQ_SYRS_778]

The PPDAQ software shall allow the host to read all the bits from all PPDIO96 boards in the system.

Note: this operation reads data from internal PPDAQ memory, it doesn't force the system to read the actual hardware bits.

Note: bits read reflect the filter type associated with each bit.

Note: if any bits are an output bit, the value read is undefined.

##### PPDAQ PPDIO96 Write Bit

[DAQ_SYRS_779]

The PPDAQ software shall allow the host to write a particular bit to the PPDIO96.

Note: bit to write is specified by (board, bank, bit).

Note: this operation writes data to internal PPDAQ memory, the value is actually written to hardware on the next polling period (every 25-100 msec).

Note: if a particular bit is programmed as an input bit, the PPDAQ software ignores the write operation.

##### PPDAQ PPDIO96 Write Bank

[DAQ_SYRS_780]

The PPDAQ software shall allow the host to write a particular bank (12 bits) on the PPDIO96.

Note: bank to write is specified by (board, bank).

Note: this operation writes data to internal PPDAQ memory, the value is actually written to hardware on the next polling period (every 25-100 msec).

Note: if a particular bank is programmed as an input, the PPDAQ software ignores the write operation.

##### PPDAQ PPDIO96 Write Board

(Deprecated)

[DAQ_SYRS_781]

The PPDAQ software shall allow the host to write to a particular PPDIO96 board (96 bits).

Note: board to write is specified by (board, bank).

Note: this operation writes data to internal PPDAQ memory, the value is actually written to hardware on the next polling period (every 25msec).

Note: if a any bank on the board is programmed as an input, the PPDAQ software ignores the write operation.

##### PPDAQ PPDIO96 Write All

(Deprecated)

[DAQ_SYRS_782]

The PPDAQ software shall allow the host to write all the bits to all PPDIO96 boards in the system.

Note: this operation writes data to internal PPDAQ memory, the value is actually written to hardware on the next polling period (every 25msec).

Note: if a any bank on any of the boards is programmed as an input, the PPDAQ software ignores the write operation.

##### PPDAQ PPDIO96/PPDI-Specific Commands

[DAQ_SYRS_783]

The PPDAQ commands specific to the PPDIO96 board shall use the following syntax:

ppdio <commands> <optional arguments> \n

Note: PPDIO commands follow the requirements specified in [DAQ_SYRS_730], [DAQ_SYRS_731], and [DAQ_SYRS_732].

###### PPDIO96 Boards Command

(Deprecated)

[DAQ_SYRS_784]

The PPDIO boards command shall accept the following syntax:

ppdio boards boards

where boards is an integer in the range 0-6 specifying the number of installed PPDIO96 boards.

###### PPDIO96 Boards Response #1

(Deprecated)

[DAQ_SYRS_785]

The PPDIO boards command shall return the command line as its response if the command is valid.

###### PPDIO96 Boards Response #2

(Deprecated)

[DAQ_SYRS_786]

The PPDIO boards command shall return "Error:range:command_line" if the boards argument is outside the range 0-6.

###### PPDIO96 Boards Response #3

(Deprecated)

[DAQ_SYRS_787]

The PPDIO boards command shall return "Error:syntax:command_line" if the command is otherwise syntactically incorrect.

###### PPDIO96 Boards Command #2

(Deprecated)

[DAQ_SYRS_787.5]

The PPDIO boards command shall accept the following syntax:

ppdio boards

Note that this command has no arguments.

###### PPDIO96 Boards #2 Response

(Deprecated)

[DAQ_SYRS_787.6]

The PPDIO boards command, with no arguments, shall return the following response:

ppdio boards: x

where x is the number of PPDIO96 boards currently programmed into the system (that is, the number of boards specified by the last "ppdio boards <boards>" command that had the numeric argument).

###### PPDIO96 SLOTS Command

[DAQ_SYRS_784.1]

The PPDIO slots command shall accept the following syntax:

ppdio slots x x x x x x

where x is either 0 (no board installed), 1 (ppdio96BoardType_c), or 2 (ppdi48BoardType_c) and indicates whether a PPDIO96 or PPDI48 board is present at board select 1, 2, …, 6 (BS1-BS6, respectively).

###### PPDIO96 SLOTS Response #1

[DAQ_SYRS_785.1]

The PPDIO slots command shall return the command line as its response if the command is valid.

###### PPDIO96 SLOTS Response #2

[DAQ_SYRS_786.1]

The PPDIO slots command shall return "Error:range:command_line" if any of the x arguments are outside the range 0-1.

###### PPDIO96 SLOTS Response #3

[DAQ_SYRS_787.2]

The PPDIO slots command shall return "Error:syntax:command_line" if there are not six numeric arguments after the slots keyword or the command is otherwise incorrect.

###### PPDIO96 SLOTS Command #2

[DAQ_SYRS_787.3]

The PPDIO slots command shall accept the following syntax:

ppdio slots

###### PPDIO96 SLOTS Response #4

[DAQ_SYRS_787.4]

The “PPDIO slots” command, with no arguments, shall return the following response:

ppdio slots x x x x x x

where x is 0 or 1 depending on whether a PPDIO96 board was registered in that slot.

###### PPDIO96 DIR Command

[DAQ_SYRS_788]

The PPDIO data direction (dir) command shall accept the following syntax:

ppdio dir board bank io

where board is a board number in the range 1..6, bank is a bank number in the range 0..7, and io is either 0 (inputs) or 1 (outputs).

###### PPDIO96 Default Data Direction

[DAQ_SYRS_789]

The system shall default to all input bits on the PPDIO96.

###### PPDIO96 DIR Response #1

[DAQ_SYRS_790]

If the "ppdio dir board bank io" command is syntactically correct, then the software shall return the command line as the response.

###### PPDIO96 DIR Response #2

[DAQ_SYRS_791]

If the "ppdio dir board bank io" contains a syntax error (other than a value out of range), then the software shall return the following response:

Error:syntax:command line \n

###### PPDIO96 DIR Response #3

[DAQ_SYRS_792]

If the "ppdio dir board bank io" contains a board value that is out of range (0 or a value greater than the number of specified boards), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 DIR Response #4

[DAQ_SYRS_793]

If the "ppdio dir board bank io" contains a bank value that is out of range (not in 0..7), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 DIR Response #5

[DAQ_SYRS_794]

If the "ppdio dir board bank io" contains an io value that is out of range (not zero or one), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 DIR Command #2

[DAQ_SYRS_794.5]

The PPDIO data direction (dir) command shall accept the following syntax:

ppdio dir board bank

where board is a board number in the range 1..6 and bank is a bank number in the range 0..7.

###### PPDIO96 DIR #2 Response

[DAQ_SYRS_794.6]

If the "ppdio dir board bank" command is syntactically correct, then the software shall return "ppdio dir: io" as the response, where io is the current data direction programmed for that bank (0 or 1).

###### PPDIO96 PULLUP Command

[DAQ_SYRS_795]

The PPDIO pull-up resistor (pullup) command shall accept the following syntax:

ppdio pullup board bank bit pullup

where board is a board number in the range 1..6, bank is a bank number in the range 0..7, bit is a hexadecimal number in the range 0..B, and pulllup is either 0 (no pullup) or 1 (insert pullup resistor).

###### PPDIO96 Default Pullup Resistor State

[DAQ_SYRS_796]

The system shall default to no pull-up resistors for all PPDIO96 input bits.

###### PPDIO96 PULLUP Response #1

[DAQ_SYRS_797]

If the "ppdio pullup board bank bit pullup" command is syntactically correct, then the software shall return the command line as the response.

###### PPDIO96 PULLUP Response #2

[DAQ_SYRS_798]

If the "ppdio pullup board bank bit pullup" contains a syntax error (other than a value out of range), then the software shall return the following response:

Error:syntax:command line \n

###### PPDIO96 PULLUP Response #3

[DAQ_SYRS_799]

If the "ppdio pullup board bank bit pullup" contains a board value that is out of range (0 or a value greater than the number of specified boards), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 PULLUP Response #4

[DAQ_SYRS_800]

If the "ppdio pullup board bank bit pullup" contains a bank value that is out of range (not in 0..7), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 PULLUP Response #5

[DAQ_SYRS_801]

If the "ppdio pullup board bank bit pullup" contains a bit value that is out of range (not in 0..B), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 PULLUP Command #2

[DAQ_SYRS_801.5]

The PPDIO pull-up resistor (pullup) command shall accept the following syntax:

ppdio pullup board bank value

where board is a board number in the range 1..6, bank is a bank number in the range 0..7, and value is a hexadecimal number in the range 0..FFF.

###### PPDIO96 PULLUP #2 Response

[DAQ_SYRS_801.6]

If the "ppdio pullup board bank bit pullup" command is syntactically correct, then the software shall return the command line as the response.

###### PPDIO96 PULLUP Command #3

[DAQ_SYRS_801.7]

The PPDIO pull-up resistor (pullup) command shall accept the following syntax:

ppdio pullup board bank

where board is a board number in the range 1..6 and bank is a bank number in the range 0..7.

###### PPDIO96 PULLUP #3 Response

[DAQ_SYRS_801.8]

If the "ppdio pullup board bank " command is syntactically correct, then the software shall return "ppdio pul: xxx" as the response (where 'xxx' is the 12-bit hexadecimal representation of the pullup resistor setting).

###### PPDIO96 POLARITY Command

[DAQ_SYRS_802]

The PPDIO polarity (polarity) command shall accept the following syntax:

ppdio polarity board bank bit polarity

where board is a board number in the range 1..6, bank is a bank number in the range 0..7, bit is a hexadecimal number in the range 0..B, and polarity is either 0 (active low) or 1 (active high).

###### PPDIO96 Default Polarity

[DAQ_SYRS_803]

The system shall default to active high polarity for all input bits on the PPDIO96.

###### PPDIO96 POLARITY Response #1

[DAQ_SYRS_804]

If the "ppdio polarity board bank bit polarity" command is syntactically correct, then the software shall return the command line as the response.

###### PPDIO96 POLARITY Response #2

[DAQ_SYRS_805]

If the "ppdio pol board bank bit polarity" contains a syntax error (other than a value out of range), then the software shall return the following response:

Error:syntax:command line \n

###### PPDIO96 POLARITY Response #3

[DAQ_SYRS_806]

If the "ppdio polarity board bank bit polarity" contains a board value that is out of range (0 or a value greater than the number of specified boards), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 POLARITY Response #4

[DAQ_SYRS_807]

If the "ppdio polarity board bank bit polarity" contains a bank value that is out of range (not in 0..7), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 POLARITY Response #5

[DAQ_SYRS_808]

If the "ppdio polarity board bank bit polarity" contains a bit value that is out of range (not in 0..B), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 POLARITY Response #6

[DAQ_SYRS_809]

If the "ppdio polarity board bank bit polarity" contains a polarity value that is out of range (not zero or one), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 POLARITY Response #7

[DAQ_SYRS_810]

If the "ppdio polarity board bank bit polarity" contains a polarity value that is out of range (not zero or one), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 POLARITY Command #2

[DAQ_SYRS_810.5]

The PPDIO polarity (pol) command shall accept the following syntax:

ppdio polarity board bank polarity

where board is a board number in the range 1..6, bank is a bank number in the range 0..7, and polarity is a 12-bit hexadecimal number in the range 0..FFF.

###### PPDIO96 POLARITY  #2 Response

[DAQ_SYRS_810.6]

If the "ppdio polarity board bank polarity" command is syntactically correct, then the software shall return the command line as the response.

###### PPDIO96 POLARITY  Command #3

[DAQ_SYRS_801.7]

The PPDIO polarity (polarity) command shall accept the following syntax:

ppdio polarity board bank

where board is a board number in the range 1..6 and bank is a bank number in the range 0..7.

###### PPDIO96 POLARITY  #3 Response

[DAQ_SYRS_801.8]

If the "ppdio polarity board bank " command is syntactically correct, then the software shall return "ppdio pol: xxx" as the response (where 'xxx' is the 12-bit hexadecimal representation of the polarity setting).

###### PPDIO96 FILTER Command

[DAQ_SYRS_811]

The PPDIO filter (filter) command shall accept the following syntax:

ppdio filter board bank bit filter

where board is a board number in the range 1..6, bank is a bank number in the range 0..7, bit is a hexadecimal number in the range 0..B, and filter is one of the following values:

- Most recent

- First

- Vote

- Loser

- Debounce

###### PPDIO96 Default Filter State

[DAQ_SYRS_812]

The system shall default to the "most recent" filter state for all PPDIO96 input bits.

###### PPDIO96 FILTER Response #1

[DAQ_SYRS_813]

If the "ppdio filter board bank bit filter" command is syntactically correct, then the software shall return the command line as the response.

###### PPDIO96 FILTER Response #2

[DAQ_SYRS_814]

If the "ppdio filter board bank bit filter" contains a syntax error (other than a value out of range), then the software shall return the following response:

Error:syntax:command line \n

###### PPDIO96 FILTER Response #3

[DAQ_SYRS_815]

If the "ppdio filter board bank bit filter" contains a board value that is out of range (0 or a value greater than the number of specified boards), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 FILTER Response #4

[DAQ_SYRS_816]

If the "ppdio filter board bank bit filter" contains a bank value that is out of range (not in 0..7), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 FILTER Response #5

[DAQ_SYRS_817]

If the "ppdio filter board bank bit filter" contains a bit value that is out of range (not in 0..B), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 FILTER Response #6

[DAQ_SYRS_818]

If the "ppdio filter board bank bit filter" contains a filter value that is out of range (not in 0-5), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 FILTER Command #2

[DAQ_SYRS_818.5]

The PPDIO filter (filter) command shall accept the following syntax:

ppdio filter board bank bit

where board is a board number in the range 1..6, bank is a bank number in the range 0..7, and bit is a hexadecimal number in the range 0..B.

###### PPDIO96 FILTER #2 Response

[DAQ_SYRS_818.6]

If the "ppdio filter board bank bit " command returns the current filter value for the specified (board, bank bit):

ppdio fltr: filter

###### PPDIO96 DBNC Command

[DAQ_SYRS_819]

The PPDAQ software shall accept the following syntax:

ppdio debounce board bank bit value

where board is a board number in the range 1..6, bank is a bank number in the range 0..7, bit is a hexadecimal bit number in the range 0..0xB, and value is a debounce time in the range 0..0x27.

###### PPDIO96 DBNC Command #2

[DAQ_SYRS_819.5]

The PPDAQ software shall accept the following syntax:

ppdio debounce board bank bit

where board is a board number in the range 1..6, bank is a bank number in the range 0..7, and bit is a hexadecimal bit number in the range 0..B.

###### PPDIO96 DBNC #2 Response

[DAQ_SYRS_820]

If the "ppdio dbnc board bank bit" command is syntactically correct, then the software shall return the following line as the response:

ppdio dbnc: debounce

where "debounce" is a hexadecimal number (1-0x27, without the 0x prefix) holding the debounce count for the selected bit.

###### PPDIO96 DBNC Response #1

[DAQ_SYRS_821]

If the "ppdio debounce board bank bit value" command is syntactically correct, then the software shall return the command line as the response.

###### PPDIO96 DBNC Response #2

[DAQ_SYRS_822]

If the "ppdio debounce board bank bit value" contains a syntax error (other than a value out of range), then the software shall return the following response:

Error:syntax:command line \n

###### PPDIO96 DBNC Response #3

[DAQ_SYRS_823]

If the "ppdio debounce board bank bit polarity" contains a board value that is out of range (0 or a value greater than the number of specified boards), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 DBNC Response #4

[DAQ_SYRS_824]

If the "ppdio debounce board bank bit polarity" contains a bank value that is out of range (not in 0..7), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 DBNC Response #5

[DAQ_SYRS_825]

If the "ppdio debounce board bank bit polarity" contains a bit value that is out of range (not in 0..B), then the software shall return the following response:

Error:range:command line \n

###### PPDIO96 DIN Command #1

(Deprecated)

[DAQ_SYRS_826]

The PPDIO digital data input (din) command shall accept the following syntax:

ppdio din

to read all the bits from all PPDIO96 boards.

###### PPDIO96 DIN Command #2

[DAQ_SYRS_827]

The PPDIO digital data input (din) command shall accept the following syntax:

ppdio din board

where the board argument is a number in the range 1..6. This reads all the input bits on the specified board.

###### PPDIO96 DIN Command #3

[DAQ_SYRS_828]

The PPDIO digital data input (din) command shall accept the following syntax:

ppdio din board bank

where the board argument is a number in the range 1..6 and the bank argument is a number in the range 0..7. This reads all the 12 input bits on the specified (board, bank).

###### PPDIO96 DIN Command #4

[DAQ_SYRS_829]

The PPDIO digital data input (din) command shall accept the following syntax:

ppdio din board bank bit

where the board argument is a number in the range 1..6, the bank argument is a number in the range 0..7, and the bit argument is a hexadecimal number in the range 0..B (with no leading radix characters). This reads the bit specified by (board, bank, bit).

###### PPDIO96 DIN Response Bits #1

[DAQ_SYRS_830]

When returning digital input data in response to a PPDIO96 din command, the software shall return the bit value as appropriate for each bit's filter type (most recent, first, vote, loser, debounce).

###### PPDIO96 DIN Response Bits #2

[DAQ_SYRS_831]

When returning digital input data in response to a PPDIO96 din command, the software shall return the bit value as appropriate for each bit's polarity setting (active low/high).

Note: Active low polarity will invert the signal received on the input port, active high polarity returns the received signal as-is.

###### PPDIO96 DIN Response #1

(Deprecated)

[DAQ_SYRS_832]

In response to a "ppdio din" command, the PPDAQ software shall return the following response:

DIN: xxx0 xxx2 … xxxn-1

where xxx represents a 3-digit (12-bit) hexadecimal number and 'n' is the number of banks present across all PPDIO96 boards.

Note: If a particular bank is programmed as output bits rather than input bits, the din command will still return a value for those 12 bits; the value is undefined and should be ignored.

Note: xxx1 corresponds to bank #0 on the first board in the system; xxxn-1 corresponds to bank #7 on the last PPDIO96 board in the system.

Note: xxx represents a 12-bit value. Bit zero of this number corresponds to the L.O. bit in the associated bank of bits and bit 11 of this number corresponds to the H.O. bit in the associated bank of bits.

###### PPDIO96 DIN Response #2

[DAQ_SYRS_833]

In response to a "ppdio din board" command, the PPDAQ software shall return the following response:

ppdio din: xxx0 xxx1 xxx2 xxx3 xxx4 xxx5 xxx6 xxx7

where xxx represents a 3-digit (12-bit) hexadecimal number.

Note: If a particular bank is programmed as output bits rather than input bits, the DIN command will still return a value for those 12 bits; the value is undefined and should be ignored.

Note: xxx represents a 12-bit value. Bit zero of this number corresponds to the L.O. bit in the associated bank of bits and bit 11 of this number corresponds to the H.O. bit in the associated bank of bits.

###### PPDIO96 DIN Response #3

[DAQ_SYRS_834]

In response to a "ppdio din board bank" command, the PPDAQ software shall return the following response:

ppdio din: xxx

where xxx represents a 3-digit (12-bit) hexadecimal number.

Note: If the specified bank is programmed as output bits rather than input bits, the DIN command will still return a value for those 12 bits; the value is undefined and should be ignored.

Note: xxx represents a 12-bit value. Bit zero of this number corresponds to the L.O. bit in the associated bank of bits and bit 11 of this number corresponds to the H.O. bit in the associated bank of bits.

###### PPDIO96 DOUT Command #1

(Deprecated)

[DAQ_SYRS_835]

The PPDIO digital data output (dout) command shall accept the following syntax:

ppdio dout xxx0 xxx1 … xxxn-1

where xxx represents a 3-digit (12-bit) hexadecimal number and 'n' is the number of banks present across all PPDIO96 boards.

Note: If a particular bank is programmed as input bits rather than output bits, the dout command will ignore the bits in the input bank; however, those placeholder bits must still be present.

Note: xxx0 corresponds to bank #0 on the first board in the system; xxxn-1 corresponds to bank #7 on the last PPDIO96 board in the system.

Note: xxx represents a 12-bit value. Bit zero of this number corresponds to the L.O. bit in the associated bank of bits and bit 11 of this number corresponds to the H.O. bit in the associated bank of bits.

Note: this command always has (board count)*8 arguments (8, 16, 24, 32, 40, or 48 arguments). The argument count is how the software differentiates this command from other DOUT commands.

###### PPDIO96 DOUT Command #2

[DAQ_SYRS_836]

The PPDIO digital data output (dout) command shall accept the following syntax:

ppdio dout board  xxx0 xxx1 xxx2 xxx3 xxx4 xxx5 xxx6 xxx7

where the board argument is a number in the range 1..6. This writes all the output bits to the specified board.

Note: If a particular bank is programmed as input bits rather than output bits, the dout command will ignore the bits in the input bank; however, those placeholder bits must still be present.

Note: this command always has nine numeric arguments. The argument count is how the software differentiates this command from other DOUT commands.

###### PPDIO96 DOUT Command #3

[DAQ_SYRS_837]

The PPDIO digital data output (dout) command shall accept the following syntax:

ppdio dout board bank xxx

where the board argument is a number in the range 1..6, the bank argument is a number in the range 0..7, and xxx is a 12-bit hexadecimal value (3 hex digits). This writes the 12 output bits to the specified (board, bank).

Note: if (board,bank) is set up as an input port, the software ignores the write operation.

Note: this command always has three numeric arguments. The argument count is how the software differentiates this command from other DOUT commands.

###### PPDIO96 DOUT Command #4

[DAQ_SYRS_838]

The PPDIO digital data output (dout) command shall accept the following syntax:

ppdio dout board bank bit value

where the board argument is a number in the range 1..6, the bank argument is a number in the range 0..7, and the bit argument is a hexadecimal number in the range 0..B (with no leading radix characters), the value argument is zero or one. This command writes the specified bit to the output.

Note: if (board,bank) is set up as an input port, the software ignores the write operation.

Note: this command always has four numeric arguments. The argument count is how the software differentiates this command from other DOUT commands.

###### PPDIO96 DOUT Command #5

[DAQ_SYRS_838.5]

The PPDIO digital data output (dout) command shall accept the following syntax:

ppdio dout board bank

where the board argument is a number in the range 1..6 and the bank argument is a number in the range 0..7. This returns the 12-bit value last written to (board, bank).

Note: if (board,bank) is set up as an input port, the command returns zero bits.

Note: this command always has two numeric arguments. The argument count is how the software differentiates this command from other DOUT commands.

###### PPDIO96 DOUT Response #1

(Deprecated)

[DAQ_SYRS_839]

In response to a syntactically correct "ppdio dout xxx0 xxx1 … xxxn-1" command, the PPDAQ software shall return the command line as the acknowledgement.

###### PPDIO96 DOUT Response #2

[DAQ_SYRS_840]

(Deprecated)

In response to a  "ppdio dout xxx0 xxx1 … xxxn-1" command where one or more of the hexadecmial values are out of range (0-FFF), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

Note:  [requirement alert] technically, this is 48 separate requirements. They have been combined into a single requirement for brevity in the SyRS.

###### PPDIO96 DOUT Response #3

[DAQ_SYRS_841]

(Deprecated)

In response to a  "ppdio dout xxx0 xxx1 … xxxn-1" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

###### PPDIO96 DOUT Response #4

[DAQ_SYRS_842]

In response to a valid "ppdio dout board  xxx0 xxx1 xxx2 xxx3 xxx4 xxx5 xxx6 xxx7" command,  the PPDAQ software shall return the command line as the acknowledgement.

###### PPDIO96 DOUT Response #5

[DAQ_SYRS_843]

In response to a  "ppdio dout board xxx0 xxx1 … xxx7" command where the board value is outside the range 1..6, the software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDIO96 DOUT Response #6

[DAQ_SYRS_844]

In response to a  "ppdio dout board xxx0 xxx1 … xxx7" command where one or more of the hexadecmial values are out of range (0-FFF), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

Note:  [requirement alert] technically, this is 8 separate requirements. They have been combined into a single requirement for brevity in the SyRS.

###### PPDIO96 DOUT Response #7

[DAQ_SYRS_845]

In response to a  "ppdio dout board xxx0 xxx1 xxx2 xxx3 xxx4 xxx5 xxx6 xxx7" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

###### PPDIO96 DOUT Response #8

[DAQ_SYRS_846]

In response to a valid "ppdio dout board bank xxx" command, the PPDAQ software shall return the command line as the acknowledgement.

###### PPDIO96 DOUT Response #9

[DAQ_SYRS_847]

In response to a  "ppdio dout board bank xxx" command where the board value is outside the range 1..6, the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDIO96 DOUT Response #10

[DAQ_SYRS_848]

In response to a  "ppdio dout board bank xxx" command where the bank value is outside the range 0..7, the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDIO96 DOUT Response #11

[DAQ_SYRS_849]

In response to a  "ppdio dout board bank xxx" command where the hexadecmial output value is out of range (0-FFF), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDIO96 DOUT Response #12

[DAQ_SYRS_850]

In response to a  "ppdio dout board bank xxx" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

###### PPDIO96 DOUT Response #13

[DAQ_SYRS_851]

In response to a valid "ppdio dout board bank bit value" command, the PPDAQ software shall return the command line as the acknowledgement.

###### PPDIO96 DOUT Response #14

[DAQ_SYRS_852]

In response to a  "ppdio dout board bank bit value" command where the board value is outside the range 1..6, the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDIO96 DOUT Response #15

[DAQ_SYRS_853]

In response to a  "ppdio dout board bank bit value" command where the bank value is outside the range 0..7, the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDIO96 DOUT Response #15

[DAQ_SYRS_854]

In response to a  "ppdio dout board bank bit value" command where the bit value is outside the range 0..B, the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDIO96 DOUT Response #16

[DAQ_SYRS_855]

In response to a  "ppdio dout board bank bit value" command where the hexadecmial output value (bit) is out of range (0-1), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDIO96 DOUT Response #17

[DAQ_SYRS_856]

In response to a  "ppdio dout board bank bit value" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

###### PPDIO96 DOUT Response #18

[DAQ_SYRS_856.5]

In response to a  "ppdio dout board bank " command the software shall return the following response:

ppdio dout: xxx \n

where "xxx" represents the last 12 bits written to the specified bank.

###### PPDIO96 Mal-Formed Command

[DAQ_SYRS_857]

If a "ppdio" command is not handled by any of the other requirements, the software shall respond with the following error acknowledgement:

Error:syntax:command_line \n

#### PPDO Serial Output Board Support (PPSSR-16 & PPRELAY-12)

To the PPDAQ software, PPSSR-16 and PPRELAY-12 interface boards both look identical (they both have 16 outputs, though four outputs on the PPRELAY-12 board are routed to an auxiliary connector rather than to relays). Therefore, the following commands apply to both types of output boards in the system.

##### PPDO Board Count

[DAQ_SYRS_900]

The PPDAQ software shall support up to 10 PPSSR-16 and/or PPRELAY-12 boards.

Note: the limit of 10 is a hardware requirement based on TTL fanout.

##### PPDO Set Board Count

[DAQ_SYRS_901]

The PPDAQ software shall allow the host computer to specify the number of serial output boards attached to the system.

##### PPDO Serial Output Bit Count

[DAQ_SYRS_902]

The PPDAQ software shall allocate 16 bits in a global "DOUT" array for each serial output board (PPSSR-16 or PPRELAY-12) connected to the system.

##### PPDO Serial Output Bit Assignments

[DAQ_SYRS_903]

The PPDAQ software shall associate DOUT bits 0-15 with the serial output board connected directly to the DAQIF board; it shall associate bits 16-31 with the serial output board daisy-chained to the first board; it shall associate bits 32-47 with the third board in the daisy-chain; etc.

##### PPDO Serial Output by Board/Bit

[DAQ_SYRS_904]

The PPDAQ software shall address individual output bits using a (board,bit) addressing scheme.

Note: board #1 is the serial output board connected directly to the DAQIF board.

Note: bit value is always a hexadecimal value in the range 0-F.

##### PPDO Serial Output by Board

[DAQ_SYRS_905]

The PPDAQ software shall allow writing all 16 bits to a specific board in one operation.

##### PPDO Serial Output Stored in Memory

[DAQ_SYRS_906]

The PPDAQ software shall store all output data into internal memory on receipt of a serial output data write operation.

Note: the write operation does not directly write the data to the serial output shift registers.

##### PPDO Serial Output Written to Hardware

[DAQ_SYRS_907]

The PPDAQ software shall write the output data to the output hardware on the 25-100 msec polling period.

##### PPDO Serial Output Data Always Written

[DAQ_SYRS_908]

The PPDAQ software shall write the output data to the output hardware every 25-100 msec even if it has not been changed by the host since the last output cycle.

##### PPDO Reading Serial Output Data

[DAQ_SYRS_909]

The PPDAQ software shall provide the abilty to read the 16 bits currently being written to a particular serial output board.

Comment: this is useful in a multi-client environment where the bits on a given board may have been written by different clients.

Note: do not assume this operation is thread-safe. A different client could rewrite the output bits immediately after the current client reads them. Multiple clients connecting to the DAQIF must handle synchronization external to the PPDAQ software.

##### PPDO-Specific Commands

[DAQ_SYRS_910]

The PPDAQ commands specific to the serial output boards shall use the following syntax:

ppdo <commands> <optional arguments> \n

Note: PPDO commands follow the requirements specified in [DAQ_SYRS_730], [DAQ_SYRS_731], and [DAQ_SYRS_732].

###### PPDO BOARDS Command

[DAQ_SYRS_911]

The "ppdo boards" command shall have the following syntax:

ppdout boards  boards

where boards is a hexadecimal value specifying the number of PPSSR-16 and PPRELAY-12 boards (combined) connected to the DAQIF board.

###### PPDO BOARDS Command Response #1

[DAQ_SYRS_912]

A valid "ppdo boards" command shall return the command line as its acknowledgement.

###### PPDO BOARDS Command Response #2

[DAQ_SYRS_913]

If the "ppdo boards" boards argument is out of range (1-10) then the software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDO BOARDS Command Response #3

[DAQ_SYRS_914]

If the "ppdo boards" command is mal-formed, then the software shall return the following error acknowledgement:

Error:syntax:command_line \n

##### PPDO DOUT Command #1

[DAQ_SYRS_915]

The PPDO digital data output (dout) command shall accept the following syntax:

ppdo dout board xxxx

where board is a serial output board number (board #1 is the board connected directly to the DAQIF board) and xxxx represents a 4-digit (16-bit) hexadecimal number.

Note: xxxx represents a 16-bit value. Bit zero of this number corresponds to the L.O. bit in the associated bits on the specified board and bit 15 (0xF) of this number corresponds to the H.O. bit on the associated board.

##### PPDO DOUT Command #2

[DAQ_SYRS_916]

The PPDO digital data output (dout) command shall accept the following syntax:

ppdo dout board bit x

where board is a serial output board number (board #1 is the board connected directly to the DAQIF board), bit is a bit index on that board (0-F),  and x is either zero or one.

Note: the PPDAQ software will diffierentiate the two dout commands based on the number of arguments.

###### PPDO DOUT Response #1

[DAQ_SYRS_917]

In response to a syntactically correct "ppdo dout board xxxx" command, the PPDAQ software shall return the command line as the acknowledgement.

###### PPDO DOUT Response #2

[DAQ_SYRS_918]

In response to a  "ppdo dout board xxxx" command where the board value is out of range (1-#boards), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDO DOUT Response #3

[DAQ_SYRS_919]

In response to a  "ppdo dout board xxxx" command where the hexadecmial xxxx value is out of range (0-FFFF), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDO DOUT Response #4

[DAQ_SYRS_920]

In response to a  "ppdo dout board xxxx" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

##### PPDO DIN Command #1

[DAQ_SYRS_921]

The PPDO digital data input (din) command shall accept the following syntax:

ppdo dout board

where board is a serial output board number (board #1 is the board connected directly to the DAQIF board).

###### PPDO DIN Response #1

[DAQ_SYRS_922]

In response to a valid  "ppdo din board " command the PPDAQ software shall return the following response:

ppdo din board bit \n

where bit is the last value (0 or 1) written to that particular (board, bit) output bit.

###### PPDO DIN Response #2

[DAQ_SYRS_923]

In response to a  "ppdo din board bit" command where the hexadecmial board value is out of range (1-10) or the bit value is out of range (0-F), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDO DIN Response #3

[DAQ_SYRS_924]

In response to a  "ppdo din board bit" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

##### PPDO DIN Command #2

[DAQ_SYRS_925]

The PPDO digital data input (din) command shall accept the following syntax:

ppdo din board

where board is a serial output board number (board #0 is the board connected directly to the DAQIF board).

Note: the PPDAQ software will diffierentiate the two "ppdo din" commands based on the number of arguments.

###### PPDO DIN Response #4

[DAQ_SYRS_926]

In response to a valid  "ppdo din board " command the PPDAQ software shall return the following response:

ppdo din board bit \n

where bit is the last value (0 or 1) written to that particular (board, bit) output bit.

###### PPDO DIN Response #5

[DAQ_SYRS_927]

In response to a  "ppdo din board bit" command where the hexadecmial board value is out of range (1-10) or the bit value is out of range (0-F), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPDO DOUT Response #6

[DAQ_SYRS_928]

In response to a  "ppdo dout board bit" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

##### PPDO TYPE Command #1

[DAQ_SYRS_929]

The PPDO type (type) command shall accept the following syntax:

ppdo type board

where board is a serial output board number (board #1 is the board connected directly to the DAQIF board).

###### PPDO TYPE Response #1

[DAQ_SYRS_930]

In response to a valid  "ppdo type board " command the PPDAQ software shall return the following response:

ppdo type: type \n

where type is 0 (no board installed), 1 (PPRELAY-12 board installed), or 2 (PPSSR-16 board installed).

###### PPDO TYPE Response #2

[DAQ_SYRS_931]

In response to a  "ppdo type board" command where the board number is outside the range 1-10, the software shall return the following response:

Error:range:command_line \n

###### PPDO TYPE Response #3

[DAQ_SYRS_932]

In response to a  "ppdo type board" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

##### PPDO TYPE Command #2

[DAQ_SYRS_933]

The PPDO type (type) command shall accept the following syntax:

ppdo type board type

where board is a serial output board number (board #0 is the board connected directly to the DAQIF board) and type is 1 or 2 (1 = PPRELAY-12 board, 2=PPSSR-16 board).

Note: the PPDAQ software will diffierentiate the two "ppdo din" commands based on the number of arguments.

###### PPDO TYPE Response #4

[DAQ_SYRS_934]

In response to a valid  "ppdo type board type " command the PPDAQ software shall return the command line as the acknowledgement.

###### PPDO TYPE Response #5

[DAQ_SYRS_935]

In response to a  "ppdo type board type" command where the board number is outside the range 1-10, the software shall return the following response:

Error:range:command_line \n

###### PPDO TYPE Response #6

[DAQ_SYRS_936]

In response to a  "ppdo type board type" command where the type number is outside the range 1-2, the software shall return the following response:

Error:range:command_line \n

###### PPDO TYPE Response #7

[DAQ_SYRS_937]

In response to a  "ppdo type boar typed" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

#### PPAIO-16/4 Support PPAIO analog data input (ain)

##### PPAIO-16/4 Board Count

[DAQ_SYRS_950]

The PPDAQ software shall support 0 to 8 PPAIO-16/4 boards.

##### PPAIO-16/4 Double-ended Input Support

[DAQ_SYRS_951]

The PPDAQ software shall allow the programming of sequential even-odd pairs of analog inputs on the PPAIO-16/4 as double-ended inputs.

##### PPAIO-16/4 Single-ended Input is Default

[DAQ_SYRS_952]

The PPDAQ software shall default all analog inputs on the PPAIO-16/4 as single-ended inputs.

##### PPAIO-16/4 Set ADC Port Count

[DAQ_SYRS_953]

The PPDAQ software shall allow the host computer to specify the number of PPAIO-16/4 ADC ports attached to the system.

##### PPAIO-16/4 Board Positions

[DAQ_SYRS_954]

PPAIO-16/4 boards shall occupy the following I2C positions:

[DAQ_SYRS_955]

- I2C 0/1: Board #0, ADC ports 0-15, DAC ports 0-3

[DAQ_SYRS_956]

- I2C 0/1: Board #4 (daisy-chained to board #0), ADC ports 64-79, DAC ports 16-19

[DAQ_SYRS_957]

- I2C 2/3: Board #1, ADC ports 16-31, DAC ports 4-7

[DAQ_SYRS_958]

- I2C 2/3: Board #5 (daisy-chained to board #1), ADC ports 80-95, DAC ports 20-23

[DAQ_SYRS_959]

- I2C 4/5: Board #2, ADC ports 32-47, DAC ports 8-11

[DAQ_SYRS_960]

- I2C 4/5: Board #6 (daisy-chained to board #2), ADC ports  96-111, DAC ports 24-27

[DAQ_SYRS_961]

- I2C 6/7: Board #3, ADC ports 48-63, DAC ports 12-16

[DAQ_SYRS_962]

- I2C 6/7: Board #7 (daisy-chained to board #3), ADC ports 112-127, DAC ports 28-31

##### PPAIO-16/4 Set Single/Double-Ended

[DAQ_SYRS_963]

The PPDAQ software shall allow the host computer to specify which output ports on a particular PPAIO-16/5 are single-ended or double-ended.

##### PPAIO-16/4 Input Buffering

[DAQ_SYRS_964]

The PPDAQ software shall preserve a minimum of 40 analog input readings (in a buffer) between host read requests.

##### PPAIO-16/4 Input FIFO Clear

[DAQ_SYRS_965]

The PPDAQ software shall clear (empty) the PPAIO-16/4 FIFO history list after the host computer reads the data.

##### PPAIO-16/4 Input FIFO Overflow

[DAQ_SYRS_966]

The PPDAQ software shall only keep the last 40 readings in the FIFO queue while waiting for the host to read the data.  If more than 40 readings are taken while waiting for a host request, the PPDAQ software shall delete the oldest (greater than 40 readings) entries.

##### PPAIO-16/4 Input FIFO Filtering

[DAQ_SYRS_967]

The PPDAQ software shall allow the following filtering options on PPAIO-16/4 input data:

[DAQ_SYRS_968]

- Most recent reading in history list (last value read from the PPAIO-16/4 )

[DAQ_SYRS_969]

- First reading in history list (first element read after the last host data request)

[DAQ_SYRS_970]

- Maximum: largest reading in history list

[DAQ_SYRS_971]

- Minimum: smallest reading in history list

[DAQ_SYRS_972]

- Mean: the arithmetic mean of all the values since the last host reading

[DAQ_SYRS_973]

- Median: the median of all the values since the last host reading

##### PPAIO-16/4 Input Setting Filtering

[DAQ_SYRS_974]

The PPDAQ software shall allow the host system to set the input filtering for individual ADC ports.

Note: filtering is specified by (board, bank, bit, filter_type) where filter type is

- 0: Most recent reading

- 1: First reading in history list

- 2: Maximum

- 3: Minimum

- 4: Mean

- 5: Median

###### PPAIO-16/4 Default Filter State

[DAQ_SYRS_975]

The system shall default to the "most recent reading" filter state for all PPAIO-16/4 inputs.

##### PPAIO-16/4 Set Gain

[DAQ_SYRS_976]

The PPDAQ software shall allow the host to set an independent gain for each PPAIO-16/4 ADC channel.

##### PPAIO-16/4 ±6.144V Gain

[DAQ_SYRS_977]

The PPDAQ software shall allow the host to set a gain setting accepting 6.144V input for the ADC channel.

##### PPAIO-16/4 ±4.096V Gain

[DAQ_SYRS_978]

The PPDAQ software shall allow the host to set a gain setting accepting ±4.096V input for the ADC channel.

Note: The PPAC-4 module will convert ±10V to ±4.096V.

##### PPAIO-16/4 ±2.048V Gain

[DAQ_SYRS_979]

The PPDAQ software shall allow the host to set a gain setting accepting ±2.048V input for the ADC channel.

Note: The PPAC-4 module will convert ±5V to ±2.048V.

##### PPAIO-16/4 ±1.024V Gain

[DAQ_SYRS_980]

The PPDAQ software shall allow the host to set a gain setting accepting ±1.024V input for the ADC channel.

Note: The PPAC-4 module will convert ±2.5V to ±1.024V.

##### PPAIO-16/4 ±0.512V Gain

[DAQ_SYRS_981]

The PPDAQ software shall allow the host to set a gain setting accepting ±0.512 input for the ADC channel.

Note: The PPAC-4 module will convert ±1.25V to ±0.512V.

##### PPAIO-16/4 Read ADC

[DAQ_SYRS_982]

The PPDAQ software shall allow the host to read a particular ADC value from the PPAIO-16/4.

Note: bit to read is specified by (board, port).

Note: this operation reads data from internal PPDAQ memory, it doesn't force the system to read the actual hardware ADCs.

Note: values read reflect the filter type associated with each ADC port.

Note: if a port is defined as double-ended and you read the odd port # of the even/odd port number pair, the result is undefined.

##### PPAIO-16/4 Read Board

[DAQ_SYRS_983]

The PPDAQ software shall allow the host to read a particular board (16 words) from the PPAIO-16/4.

##### PPAIO-16/4 Write DAC Port

[DAQ_SYRS_984]

The PPDAQ software shall allow the host to write a particular DAC port on the PPAIO-16/4.

Note: DAC to write is specified by (board, port) where board=1..8 and port is 0-3.

##### PPAIO-Specific Commands

[DAQ_SYRS_985]

The PPDAQ commands specific to the serial output boards shall use the following syntax:

ppaio <commands> <optional arguments> \n

Note: PPAIO commands follow the requirements specified in [DAQ_SYRS_730], [DAQ_SYRS_731], and [DAQ_SYRS_732].

###### PPAIO Boards Command

[DAQ_SYRS_986]

The PPAIO boards command shall accept the following syntax:

ppaio boards boards

where boards is an integer in the range 0-8 specifying the number of installed PPAIO boards.

###### PPAIO Boards Response #1

[DAQ_SYRS_987]

The PPAIO boards command shall return the command line as its response if the command is valid.

###### PPAIO Boards Response #2

[DAQ_SYRS_988]

The PPAIO boards command shall return "Error:range:command_line" if the boards argument is out of the range 1..8.

###### PPAIO Boards Response #3

[DAQ_SYRS_989]

The PPAIO boards command shall return "Error:syntax:command_line" if the command is otherwise syntactically incorrect.

###### PPAIO Gain Command

[DAQ_SYRS_990]

The PPAIO gain command shall accept the following syntax:

ppaio gain board  port  gain

where board is an integer in the range 1..8 specifying the specific PPAIO board, port is a hexadecimal value in the range 0-F specifying the particular port on the board, and gain is a value in the range 0-4 representing one of the following gain adjustments:

0:	±6.144V

1:	±4.096V

2:	±2.048V

3:	±1.024V

4:	±0.512V

###### PPAIO Default Gain

[DAQ_SYRS_991]

Until otherwise reprogrammed, the PPDAQ software shall use a gain of 1 (4.096V) as the default gain for all ADC channels.

###### PPAIO Gain Response #1

[DAQ_SYRS_992]

If the "ppaio gain board  port  gain" command is syntactically correct, then the software shall return the command line as the response.

###### PPAIO Gain Response #2

[DAQ_SYRS_993]

If the "ppaio gain board  port  gain" contains a syntax error (other than a value out of range), then the software shall return the following response:

Error:syntax:command line \n

###### PPAIO Gain Response #3

[DAQ_SYRS_994]

If the "ppaio gain board  port  gain" contains a board value that is out of range (0 or a value greater than the number of specified boards), then the software shall return the following response:

Error:range:command line \n

###### PPAIO Gain Response #4

[DAQ_SYRS_995]

If the "ppaio gain board  port  gain" contains a port value that is out of range (not in 0-F), then the software shall return the following response:

Error:range:command line \n

###### PPAIO Gain Response #5

[DAQ_SYRS_996]

If the "ppaio gain board  port  gain" contains a port value that odd and is part of a double-ended analog input channel, then the software shall return the following response:

Error:range:command line \n

###### PPAIO Gain Response #6

[DAQ_SYRS_997]

If the "ppaio gain board  port  gain" contains a gain value that is out of range (not in 0-4), then the software shall return the following response:

Error:range:command line \n

###### PPAIO FILTER Command

[DAQ_SYRS_998]

The PPAIO filter (fltr) command shall accept the following syntax:

ppaio FILTER board port filter

where board is a board number in the range 1..8, port is a port number in the range 0-F, and filter is one of the following values:

- 0: Most recent reading

- 1: First reading in history list

- 2: Maximum

- 3: Minimum

- 4: Mean

- 5: Median

###### PPAIO FILTER Response #1

[DAQ_SYRS_999]

If the "ppaio FILTER board port filter" command is syntactically correct, then the software shall return the command line as the response.

###### PPAIO FILTER Response #2

[DAQ_SYRS_1000]

If the "ppaio FILTER board port filter" contains a syntax error (other than a value out of range), then the software shall return the following response:

Error:syntax:command line \n

###### PPAIO FILTER Response #3

[DAQ_SYRS_1001]

If the "ppaio FILTER board port filter" contains a board value that is out of range (0 or a value greater than the number of specified boards), then the software shall return the following response:

Error:range:command line \n

###### PPAIO FILTER Response #4

[DAQ_SYRS_1002]

If the "ppaio FILTER board port filter" contains a port value that is out of range (not in 0-F), then the software shall return the following response:

Error:range:command line \n

###### PPAIO FILTER Response #5

[DAQ_SYRS_1003]

If the "ppaio FILTER board port filter" contains a port value that odd and is part of a double-ended analog input channel, then the software shall return the following response:

Error:range:command line \n

###### PPAIO FILTER Response #6

[DAQ_SYRS_1004]

If the "ppaio FILTER board port filter" contains a gain value that is out of range (not in 0..5), then the software shall return the following response:

Error:range:command line \n

###### PPAIO AIN Command #1

[DAQ_SYRS_1005]

(Deprecated)

The PPAIO analog data input (ain) command shall accept the following syntax:

ppaio ain

to read all the values from all (initialized) PPAIO-16/4 boards.

###### PPAIO AIN Command #2

[DAQ_SYRS_1006]

The PPAIO analog data input (ain) command shall accept the following syntax:

ppaio ain board

where the board argument is a number in the range 1..8. This reads all the input ports on the specified board.

###### PPAIO AIN Command #3

[DAQ_SYRS_1007]

The PPAIO analog data input (ain) command shall accept the following syntax:

ppaio ain board port

where the board argument is a number in the range 1..8 and port is a number in the range 0..7. This the specified input port.

###### PPAIO AIN Response Value

[DAQ_SYRS_1008]

When returning digital input data in response to a PPAIO ain command, the software shall return the ADC port value (16-bit hexadecimal value representing a two's complement signed number) as appropriate for each bit's filter type (most recent, first, minimum, maximum, mean, or median).

###### PPAIO AIN Response #1

(Deprecated)

[DAQ_SYRS_1009]

In response to a "ppaio ain " command, the PPDAQ software shall return the following response:

AIN: xxxx0 xxxx2 … xxxxn-1

where xxxx represents a 4-digit (16-bit) hexadecimal number and 'n' is the number of ports present across all PPAIO-16/4 boards.

Note: xxxx1 corresponds to port #0 on the first board in the system; xxxxn-1 corresponds to port #15 on the last PPAIO-16/4 board in the system.

Note: Values corresponding to odd port numbers of a double-ended input pair are undefined.

###### PPAIO AIN Response #2

[DAQ_SYRS_1010]

In response to a "ppaio ain board" command, the PPDAQ software shall return the following response:

AIN: xxxx0 xxxx2 … xxxx15

where xxxx represents a 4-digit (16-bit) hexadecimal number.

Note: Values corresponding to odd port numbers of a double-ended input pair are undefined.

###### PPAIO AIN Response #3

[DAQ_SYRS_1011]

In response to a "ppaio ain board port" command, the PPDAQ software shall return the following response:

AIN: xxxx

where xxxx represents a 4-digit (16-bit) hexadecimal number.

Note: Values corresponding to odd port numbers of a double-ended input pair are undefined.

###### PPAIO AOUT Command #1

(Deprecated)

[DAQ_SYRS_1012]

The PPAIO analog data output (aout) command shall accept the following syntax:

ppaio aout xxxx0 xxxx1 … xxxxn-1

where xxx represents a 4-digit (16-bit) hexadecimal number and 'n' is the number of ports present across all PPAIO-16/4 boards.

Note: xxxx0 corresponds to port #0 on the first board in the system; xxxn-1 corresponds to port #15 on the last PPAIO-16/4 board in the system.

Note: this command always has (board count)*16  arguments (16, 32, 48, 64, 80, 96, 112, or 128 arguments). The argument count is how the software differentiates this command from other AOUT commands.

###### PPAIO AOUT Command #2

[DAQ_SYRS_1013]

The PPAIO analog data output (aout) command shall accept the following syntax:

ppaio aout board xxxx0 xxxx1 xxxx2 xxxx3

where the board argument is a number in the range 0..7. This writes all the output bits to the specified board.

Note: this command always has five numeric arguments. The argument count is how the software differentiates this command from other AOUT commands.

###### PPAIO AOUT Command #3

[DAQ_SYRS_1014]

The PPAIO analog data output (aout)  command shall accept the following syntax:

ppaio aout board port xxxx

where the board argument is a number in the range 0..7, the port argument is a number in the range 0..3, and xxxx is a 16-bit hexadecimal value (4 hex digits). This writes the 4 output values to the specified (board, port).

Note: this command always has three numeric arguments. The argument count is how the software differentiates this command from other AOUT commands.

###### PPAIO AOUT Response #1

(Deprecated)

[DAQ_SYRS_1015]

In response to a syntactically correct "ppaio aout xxx0 xxx1 … xxxn-1" command, the PPDAQ software shall return the command line as the acknowledgement.

###### PPAIO AOUT Response #2

(Deprecated)

[DAQ_SYRS_1016]

In response to a  "ppaio aout xxx0 xxx1 … xxxn-1" command where one or more of the hexadecmial values are out of range (0-FFFF), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

Note:  [requirement alert] technically, this is 16 separate requirements. They have been combined into a single requirement for brevity in the SyRS.

###### PPAIO AOUT Response #3

(Deprecated)

[DAQ_SYRS_1017]

In response to a  "ppaio aout xxxx0 xxxx1 … xxxxn-1" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

###### PPAIO AOUT Response #4

[DAQ_SYRS_1018]

In response to a valid "ppaio aout board  xxx0 xxx1 xxx2 xxx3" command,  the PPDAQ software shall return the command line as the acknowledgement.

###### PPAIO AOUT Response #5

[DAQ_SYRS_1019]

In response to a  "ppaio aout board xxx0 xxx1 … xxx3" command where the board value is outside the range 0..7, the software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPAIO AOUT Response #6

[DAQ_SYRS_1020]

In response to a  "ppaio aout board xxx0 xxx1 … xxx3" command where one or more of the hexadecmial values are out of range (0-FFFF), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

Note:  [requirement alert] technically, this is 4 separate requirements. They have been combined into a single requirement for brevity in the SyRS.

###### PPAIO AOUT Response #7

[DAQ_SYRS_1021]

In response to a  "ppaio aout board xxx0 xxx1 xxx2 xxx3" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

###### PPAIO AOUT Response #8

[DAQ_SYRS_1022]

In response to a valid "ppaio aout board port xxxx" command, the PPDAQ software shall return the command line as the acknowledgement.

###### PPAIO AOUT Response #9

[DAQ_SYRS_1023]

In response to a  "ppaio aout board port xxxx" command where the board value is outside the range 1..8, the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPAIO AOUT Response #10

[DAQ_SYRS_1024]

In response to a  "ppaio aout board port xxxx" command where the port value is outside the range 0..F, the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPAIO AOUT Response #11

[DAQ_SYRS_1025]

In response to a  "ppaio aout board port xxxx" command where the hexadecmial output value is out of range (0-FFF), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPAIO AOUT Response #12

[DAQ_SYRS_1026]

In response to a  "ppaio aout board port xxxx" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

###### PPAIO AOUT Response #13

[DAQ_SYRS_1027]

In response to a  "ppaio aout board port value" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

###### PPAIO TYPE Command #1

[DAQ_SYRS_1028]

The PPAIO type  command shall accept the following syntax:

ppaio type board port se_de

where the board argument is a number in the range 0..7, the port argument is a number in the range 0..F, and se_de is either zero (single-ended. This writes the 4 output values to the specified (board, port).

###### PPAIO TYPE Response #1

[DAQ_SYRS_1029]

In response to a  "ppaio type board port se_de" command where the hexadecmial board value is out of range (0-7), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPAIO TYPE Response #2

[DAQ_SYRS_1030]

In response to a  "ppaio type board port se_de" command where the hexadecmial port value is out of range (0-F), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPAIO TYPE Response #3

[DAQ_SYRS_1031]

In response to a  "ppaio type board port se_de" command where the hexadecmial se_de value is out of range (0-1), the PPDAQ software shall return the following error acknowledgement:

Error:range:command_line \n

###### PPAIO AOUT Response #4

[DAQ_SYRS_1032]

In response to a  "ppaio type board port se_de" command where the command is mal-formed, the software shall return the following error acknowledgement:

Error:syntax:command_line \n

###### PPAIO AOUT Response #5

[DAQ_SYRS_1033]

In response to a  "ppaio type board port se_de" command where the command is valid, the software shall return the following error acknowledgement:

ppaio type: x  \n

where 'x' is '0' or '1' representing single-ended operation or double-ended operation, respectively.

###### PPAIO Mal-Formed Command

[DAQ_SYRS_1034]

If a "ppaio" command is not handled by any of the other requirements, the software shall respond with the following error acknowledgement:

Error:syntax:command_line \n

## System Security

[DAQ_SYRS_1100]

System security shall be maintained by limiting access to the Ethernet, serial, and USB connections on the PPDAQ.

## System Operations

### Human Factors

#### Boot and Shutdown

[DAQ_SYRS_1200]

The PPDAQ shall automatically begin operation upon application of power.

#### Module Reset and Initialization

[DAQ_SYRS_1201]

The system shall reset and initialize each each attached I/O module after a power interruption.

## System Power

### PPDAQ Input Power

[DAQ_SYRS_1300]

DC power, +5V digital, shall be provided to the PPDAQ (Netburner Interface) module.

### PPAC4 Input Power

[DAQ_SYRS_1301]

DC power, ±12V clean analog, shall be provided to the PPAC4 (four-channel analog conditioning) module.

### PPAIO16-4 Input Power

[DAQ_SYRS_1302]

DC power, +5V clean analog, shall be provided to the PPAIO16-5 (16-channel analog input, 4-channel analog output) module.

### PPDIO96 Input Power

[DAQ_SYRS_1303]

DC power, +5V digital, shall be provided to the PPDIO96 (96-channel digital I/O) module.

### PPOPTO12 Input Power

[DAQ_SYRS_1304]

DC power, +5V digital, shall be provided to the PPOPTO12 (12-channel optoisolation) module.

### PPRELAY Input Power

[DAQ_SYRS_1305]

DC power, +5V digital, shall be provided to the PPRELAY12 (12-channel relay) module.

### PPSSR16 Input Power

[DAQ_SYRS_1306]

DC power, +5V digital, shall be provided to the PPSSR16 (16-channel solid-state relay) module.

### Netburner Input Power

[DAQ_SYRS_1307]

DC power, +9V digital, shall be provided to the Netburner module.

### Netburner Input Power -- USB

[DAQ_SYRS_1308]

The Netburner MOD-70 evaluation board shall be configured to take power from the J9 connector (power supply connector) rather than the micro-USB port.

### Ethernet Connection

[DAQ_SYRS_1309]

The system shall interface to the Ethernet using a standard RJ-45 Ethernet connection.

### RS-232 Serial Connection

[DAQ_SYRS_1310]

The system shall interface via RS-232 serial using a male DB-9/DTE connection.

### USB Connection

[DAQ_SYRS_1311]

The system shall interface via USB using a standard USB-B or Micro-USB-B connection.

# Design Goals and Other Requirements

# Footnotes

MCP23S17 GPIO pin assignments are not a requirement. They are provided here for informational purposes. The requirement is only the pin and bit number correlations.

MCP23S17 GPIO pin assignments are not a requirement. They are provided here for informational purposes. The requirement is only the pin and bit number correlations.

MCP23S17 GPIO pin assignments are not a requirement. They are provided here for informational purposes. The requirement is only the pin and bit number correlations.

MCP23S17 GPIO pin assignments are not a requirement. They are provided here for informational purposes. The requirement is only the pin and bit number correlations.

MCP23S17 GPIO pin assignments are not a requirement. They are provided here for informational purposes. The requirement is only the pin and bit number correlations.

MCP23S17 GPIO pin assignments are not a requirement. They are provided here for informational purposes. The requirement is only the pin and bit number correlations.

MCP23S17 GPIO pin assignments are not a requirement. They are provided here for informational purposes. The requirement is only the pin and bit number correlations.

MCP23S17 GPIO pin assignments are not a requirement. They are provided here for informational purposes. The requirement is only the pin and bit number correlations.